PT GALAKSI MEGATAMA INDONESIA Manufacturer of record

CERTIFICATION INFORMATION REQUIREMENTS — FILLED WORKING COPY

# IECEx/ATEX Certification Information Requirements

Gas Leak Detector (GLD) V2

This document follows, item by item, the section order of the ExCB checklist as it appears in the internal working-draft template (source file below), filled in with GLD's confirmed design data. Task-style annotations from that working draft have been rewritten here as professional status notes.

Document Control

| Document title | IECEx/ATEX Certification Information Requirements — Gas Leak Detector (GLD) V2 (Filled Working Copy) |
| --- | --- |
| Document no. | GMI/GLD/IECEX-CIR/2026-001 |
| Revision | 1.1 |
| Date | 7 October 2026 |
| Status | Working Document — Draft for Internal Review |
| Classification | Confidential — prepared for ATEX/IECEx certification body (ExCB) submission |
| Based on (source template) | IECEx ATEX Certification Information Requirements_Rev18092026.docx (internal working draft, 18 Sep 2026) |
| Applicant / Manufacturer | PT Galaksi Megatama Indonesia |
| Companion document | Dokumen_Teknis_Sertifikasi_GLD_IECEx_ATEX (exhaustive item-by-item status tracking) — authoritative for item-level status where the two documents differ |

Scope of this document. This is a compilation of design evidence and an honest readiness assessment against the ExCB checklist, prepared in support of a future submission — it is not itself a certificate and does not constitute self-certification. The product is still at from-scratch development/prototype stage; no finalized, serialized units exist yet.

Revision History

| REVISION | DATE | DESCRIPTION |
| --- | --- | --- |
| 0.1 | 22 September 2026 | Initial issue — filled in against the internal working-draft template (IECEx ATEX Certification Information Requirements_Rev18092026.docx), covering all three sections of the checklist. |
| 0.2 | 22 September 2026 | Total weight corrected to 2.3 kg; IP rating confirmed as IP66; operating humidity confirmed as 10–90% RH non-condensing; cable gland confirmed as M20×1.5 (IP66-rated) — all raised to Final (Section 2.3). Power consumption rounded to 8 W. Removed the ambiguous “no plastic or PVC” parenthetical from the product description (Section 2.1) and generalized “L-bracket” to “bracket” in that same narrative sentence (the dimensioned bracket/U-bolt specification in Section 2.3 is unaffected). “GLD V2 (Version 2)” simplified to “GLD V2”. Layout fix: kept the Section 2.6.b BOM paragraph together across the page break. |
| 0.3 | 22 September 2026 | “Bracket”/“L-bracket” terminology corrected throughout to “U-bolt mounting plate” (Sections 2.1, 2.3, 2.6.a), matching the actual CAD drawing (a flat 250×250 mm plate secured by 2× U-bolts — not an angled/L-shaped bracket). |
| 0.4 | 24 September 2026 | Total weight, operating temperature, and operating humidity (Section 2.3) updated to 2.378 kg, −20°C to +60°C, and 5–95% RH respectively, per current project specification data. |
| 0.5 | 28 September 2026 | Type of protection narrowed to a single team recommendation, Ex d (flameproof enclosure), in Section 2.5 — selected over Ex e and Ex i based on the existing ADC12 die-cast enclosure with a stainless-steel mesh flame path and the MQ-sensor heaters' continuous, unregulated power draw. Section 2.6.e updated accordingly. Still an internal recommendation, not an ExCB decision. |
| 0.6 | 28 September 2026 | Applicant/Manufacturer identity simplified to a single entity, PT Galaksi Megatama Indonesia (document number prefix changed to GMI/...). The separate “design and development authority” and “technical partner” entries were removed (letterhead, Document Control, Sections 1.2, 1.4). |
| 0.7 | 30 September 2026 | Section 2.6.f populated with a preliminary internal surface-temperature measurement (normal, maximum-load and fan-stalled conditions), extrapolated to +60°C ambient: highest value ≈114°C against an effective T4 limit of 130°C. Status raised to Partially available; open measurement items listed. |
| 0.8 | 7 October 2026 | Section 2.6.a extended with the released main-board electrical schematic (26 sheets, Section 2.6.a.7) and ten annotated PCB layout views with board data (Section 2.6.a.8). |
| 0.9 | 7 October 2026 | Section 2.6.a.1 added: complete mechanical, enclosure and assembly drawing set (ATEX CASING v2 sheet and six CAD views, all seven GLD ATEX CASE v3 sheets, BP18-1Z supplier drawing, component/structure/terminal/grounding drawings, assembly sequence and MQ sensor dimension drawings), with a drawing register. |
| 1.0 | 7 October 2026 | Full audit against the reviewer's revision guide: added 2.4.1 photo-set coverage; 2.6.b.1–2 Ex-critical mechanical (EX-01…17) and electrical (EL-01…13) BOM; 2.6.c.1 non-metallic materials register (MAT-01…11); 2.6.d.1 manufacturing process control d.1–d.13; 2.6.e.1–3 gas path, ignition-source assessment IS-01…17 and protection concept e.3.1–e.3.11; 2.6.f.1 temperature-class structure f.1–f.12 with results table; 2.6.g.1 usage and installation instructions g.2–g.20; 2.6.h.1 proposed nameplate (gas only, −20…+60 °C); 2.6.i.1 required Ex component certificates; 3.3–3.4 sample register (GLD2-0x1001) and fixtures; compliance matrix. |
| 1.1 | 7 October 2026 | Section 2.4.2 added: photo set of the current production configuration (23 photographs, Figures 4-101–4-123); EX-05 fan model (CIXIKEJI CX5010B5H, 5 V 0.23 A), EX-07 polymer prototype gland to be replaced by a certified Ex d gland, EX-08/EX-12/EX-17 and photo-coverage table updated. |

Table of Contents

Right-click and choose Update Field to generate the table of contents.

# 1. Basic Information (Application and Organization)

Application form (provided by ExCB as a template). Manufacturer's business license/company registration certificate. Manufacturer's organizational chart and contact information. Address of manufacturing plant and brief introduction of production facilities. (If applicable) ISO 9001 certificate, quality manual, and directory of procedure documents (for QAR/QAN review).

## 1.1 · Application Form

Issued by the ExCB as a template; not yet received.

Status: Template not yet received from the ExCB.

## 1.2 · Manufacturer's Business License / Company Registration Certificate

The Applicant and Manufacturer named in this application is PT Galaksi Megatama Indonesia, which holds the Indonesian industrial manufacturing business license (NIB) relevant to fabricating the equipment.

| Particular | Value |
| --- | --- |
| Registered legal name | PT Galaksi Megatama Indonesia |
| Legal form | Perseroan Terbatas (limited liability company), National Private / PMDN, Small Enterprise scale |
| Business registration number (NIB) | 0510220020593, issued 5 Oct 2022 (OSS system) |
| Taxpayer ID (NPWP) | 61.126.496.1-427.000, registered 29 Sep 2022 |
| Deed of establishment | Deed No. 52, 22 Sep 2022 (Notary Steffi Alphanie, S.H., M.Kn., Cilegon City) |
| Ministerial ratification | MOLHR Decree No. AHU-0067398.AH.01.01.Year 2022 (29 Sep 2022) |
| Registered (domicile) address | Plaza Summarecon Bekasi, Jl. Bulevar Ahmad Yani Kav. K.01, Level 7, Bekasi City, West Java 17143 |
| Relevant KBLI (business scope) | 25119, 25120 (tanks/reservoirs/containers of metal, incl. pressurized gas cylinders), 28221, 25920 |

Status: Available — legal documents on file with sworn English translations (Fikri Said Obed, S.S., 19 Sep 2026). Legal form/shareholding, business registration, and domicile address confirmed; ISO 9001 certification status addressed separately below.

## 1.3 · Manufacturer's Organizational Chart and Contact Information

The statutory Board of Directors and Board of Commissioners are known from the deed of establishment:

| Position | Name |
| --- | --- |
| President Director | Nur Rohman |
| Director | Antonius Prasetyo |
| Commissioner | Crisa Andi Sujatmiko |

Status: Partially available — statutory officers known. A functional organizational chart (design/production/quality reporting lines) and named ExCB correspondence contacts (certification project contact, technical contact, quality contact) have not yet been provided.

## 1.4 · Address of Manufacturing Plant and Brief Introduction of Production Facilities

Registered office and NIB-registered production sites of the Manufacturer, PT Galaksi Megatama Indonesia:

| Site | Address / Role |
| --- | --- |
| Registered office | Plaza Summarecon Bekasi, Jl. Bulevar Ahmad Yani Kav. K.01, Level 7, Bekasi City, West Java |
| Production site 1 | Jl. Kp. Jatipilar, Cikarang Selatan, Bekasi Regency |
| Production site 2 | Jl. WR Supratman, Mustikajaya, Bekasi City |
| Electronics design/assembly | Native EasyEDA/JLCPCB project; fabrication/assembly provider and its address to be confirmed in writing |

Status: Partially available — registered addresses on file (general metal-fabrication license, KBLI 25120 covers pressurized gas cylinders; not an Ex-specific manufacturing license). Floor area, principal equipment, staffing, and nominal throughput (“brief introduction” requested by the checklist) not yet provided. Production location for serial units has not been fixed — the product is at prototype stage.

## 1.5 · ISO 9001 Certificate, Quality Manual, and Directory of Procedure Documents (for QAR/QAN review)

Whether the manufacturer currently holds ISO 9001 certification, and if so its certificate number, scope, issuing body, and validity period, is to be confirmed. Quality manual and procedure-document index not yet provided.

Status: To be confirmed / to be provided. Under both IECEx and ATEX schemes, a quality system covering Ex production is expected before certificates are issued (IECEx Quality Assessment Report, or the corresponding production-quality/product-verification route under ATEX) whether or not ISO 9001 is held; the applicable route is to be confirmed with the ExCB.

# 2. Technical Documentation

## 2.1 · Detailed Product Description

The Gas Leak Detector (GLD) is an IoT-based, multi-sensor gas leak detection device designed for the early detection of flammable and process gases in oil & gas refinery environments (process units, tank farms, pipe racks, and storage/loading-unloading areas).

Functionally, the GLD integrates eight channels of metal-oxide semiconductor gas sensors (MQ series), an edge-AI microcontroller/processor (ESP32-S3), and a LoRa radio module (star-topology wireless transmission) within a single fixed-point unit installed at locations with gas-leak risk. An on-device AI gas-classification model runs directly on the unit so that detection decisions do not depend on a continuous connection to a central server. When gas concentration exceeds a defined threshold, the unit triggers a local alarm (an integrated visual/audible alarm module) and simultaneously transmits an alarm notification over the LoRa network to the operator dashboard.

The enclosure is designed for hazardous-area deployment at refinery sites, using metal materials (aluminum alloy — ADC12 die-cast grade — and stainless steel) and mounted via a U-bolt mounting plate to existing structures without drilling or welding. The current production power configuration is continuous 24 VDC, supplied via an AC/DC adapter (220 VAC/50 Hz → 24 VDC) connected to the site electrical supply. The Node Sensor (GLD) itself receives only 24 VDC at its terminal and never carries 220 VAC internally — the AC/DC adapter is a separate, site-supplied device outside the physical boundary of the certified unit.

Important: this design-intent description does not constitute a claim that the enclosure has passed testing or has been Ex-certified. The explosion-protection scheme, gas group, temperature class, and target installation zone are addressed in Section 2.5.

## 2.2 · Product Name, Model, and Specification List

| Field | Value |
| --- | --- |
| Product name | Gas Leak Detector (GLD) |
| Model / version | GLD V2 |
| Manufacturer | PT Galaksi Megatama Indonesia |
| Program owner | PT Pertamina Patra Niaga (initial deployment site: Refinery Unit IV, Cilacap) |
| Primary function | Acquisition of 8-channel gas sensor data and LoRa transmission |
| Microcontroller | ESP32-S3-WROOM-1U-N16R8 |
| LoRa radio module | E22-900MM22S |
| Dimensions (L×W×H) | 200 × 90 × 290 mm |
| Enclosure material | Aluminum alloy (ADC12 die-cast) + stainless steel |

## 2.3 · Complete and Clear Functional Description and Technical Parameters (Electrical, Mechanical, etc.)

Functional workflow (normal operating mode): sense → process → transmit. Each of the eight gas sensors continuously samples ambient conditions → data is normalized and classified by the on-device AI model (ESP32-S3) → the result is transmitted over the LoRa network (star-topology transmission) at a configurable interval, or immediately (event-driven) when an alarm condition is detected. Gas alarms are triggered through two parallel channels: a local visual/audible alarm module on the unit itself, and a push notification transmitted over the LoRa network to the dashboard.

Electrical Parameters:

| Parameter | Specification |
| --- | --- |
| Input voltage | 24 VDC (continuous) |
| Source adapter (if AC-fed) | 220 VAC / 50 Hz input → 24 VDC output |
| Internal voltage rails | 5 VDC (sensor/heater circuitry), 3.3 VDC (ESP32-S3 logic) |
| Measured power consumption | 8 W @ 24 VDC (≈ 0.33 A) — production configuration |
| Cable/PSU sizing margin | ≥ 1 A per unit (sizing margin only; actual draw is well below this) |
| Power cable | 2-conductor, labeled L+/L− (positive/negative of the 24 VDC supply), each conductor ≈ 0.75 mm diameter |
| Electrical protection | Resettable PPTC fuses (F1/F2), TVS/ESD suppression diodes, Schottky diodes on power rails |

Sensor and Communication Parameters:

| Parameter | Specification |
| --- | --- |
| Gas sensor array | 8 channels, metal-oxide (MQ-2, MQ-3B, MQ-4, MQ-5, MQ-6, MQ-7B, MQ-8, MQ-135) |
| Signal acquisition | 24-bit multi-channel ADC (ADS1256IDBR, 30,000 SPS), 500 ms scan interval |
| AI gas-classification model | On-device classifier — 3 classes: Clean Air, LPG, H₂; alarm at confidence ≥ 30%. Does not yet cover CO₂, Benzene, CO, or H₂S |
| Temperature/humidity telemetry | SHT40 sensor (auxiliary) |
| LoRa radio | E22-900MM22S, star-topology transmission, 920–923 MHz ISM band (Indonesia) |
| Data transmission interval | Configurable, default 10 s; alarm events transmitted immediately |
| RS-485 / Modbus interface | Read-only Modbus RTU slave, 9600 bit/s 8N1, Unit ID 1 (THVD1410DR) — 8 read-only registers for external integration; not used for product control |

Mechanical Parameters:

| Parameter | Specification | Status |
| --- | --- | --- |
| Dimensions (L×W×H) | 200 × 90 × 290 mm | Final |
| Total weight | 2.378 kg | Final (project-confirmed, not yet calibration-verified) |
| Enclosure material | Aluminum alloy (ADC12 die-cast) + stainless steel | Final (see caveat below) |
| Mounting method | U-bolt mounting plate (2″/DN50), no drilling/welding | Final |
| Ingress protection (IP rating) | IP66 | Final |
| Cable entry (gland) | M20×1.5 cable gland (IP66-rated), power cable L+/L−, ≈ 0.75 mm/conductor | Final (see note below) |
| Antenna mounting | External, SMA male connector | Final |
| Operating temperature range | −20°C to +60°C (ambient) | Final (project-confirmed) |
| Operating humidity range | 5–95% RH, non-condensing | Final (project-confirmed) |

Caveat on the ADC12 enclosure grade. ADC12 is the material called out on a Chinese enclosure-component manufacturer's drawing for a related die-cast sub-component (“Universal Base,” part no. BP18-1Z, from the same casing-partner supply chain — Section 2.6.a); the “GLD ATEX CASE v3” drawing of the main enclosure body itself does not carry its own material callout. ADC12 has been confirmed as the enclosure grade on the basis that the same casing manufacturer's die-cast product line uses this alloy consistently — an extrapolation from the supplier's sub-component drawing, not a direct material callout on the main enclosure drawing. This should be verified against the casing partner's own datasheet for the enclosure body before formal ExCB submission.

Caveat on cable gland. M20×1.5 matches the thread size used on the BP18-1Z reference drawing from the same casing manufacturer (Section 2.6.a) and is a standard, IP66-rated gland size well suited to a 2-conductor cable with ≈0.75 mm conductors — selected by engineering judgment from that reference, not from an independent gland datasheet or fit test on GLD's own enclosure. This should be confirmed by test/fit-check before formal ExCB submission. The operating temperature and humidity ranges above are current project-specification values (24 Sep 2026), not yet documented against a calibrated test record.

Three operating modes are distinguished at firmware level:

| Mode | Purpose | Use restriction |
| --- | --- | --- |
| Inference | Normal gas-class inference and alarm operation | Production operating mode after commissioning |
| Dataset | Engineering / data-collection activity | Not a substitute for approved detection operation |
| Nulling | Sensor baseline / calibration activity | Clean air only; approved procedure; requires complete 8/8 result |

## 2.4 · Clear Product Photos (Overall and Key Components)

The photographs below were taken directly from the Node Sensor (GLD) V2 prototype unit.

Assembled unit — motherboard, sensor module, mesh cover, casing, external antenna, and local alarm module.

Enclosure, rear view — mounting flange and cable-entry area.

#### 2.4.2 Current production configuration — assembly photo set (7 October 2026)

Photographs of one unit of the current production configuration (latest enclosure and main board), taken in assembly order from the empty base enclosure to the closed unit. No Ex marking is shown; the MQ sensor mesh is not described as flame-arresting.

| Figure 4-101. GLD V2 – Base enclosure before assembly. Machined bore of the base enclosure with the PCB mounting bosses; the external antenna is fitted at the top and its internal coaxial lead (U.FL) is routed inside. | Figure 4-102. GLD V2 – Main board placed in the base enclosure. Main PCB with the ESP32-S3 module and LoRa module seated on the mounting bosses. |
| --- | --- |
| Figure 4-103. GLD V2 – Antenna lead connected to the main board. U.FL coaxial connector of the antenna lead plugged onto the main board. | Figure 4-104. GLD V2 – PCB cover / terminal board fitted. Terminal board with screw terminals labelled RS485 A/B, FAN 5V/0V, ALARM, BAT −/+ and 24V +/−; eight 2×4 headers for the sensor modules and a micro-USB service port. |
| Figure 4-105. GLD V2 – Sensor module. MQ-series sensor with its integrated stainless-steel protective mesh, mounted on a round carrier board. | Figure 4-106. GLD V2 – Sensor module being inserted. Sensor module plugged into one of the eight headers. |
| Figure 4-107. GLD V2 – Sensor modules fitted. Sensor modules installed on the terminal board positions. | Figure 4-108. GLD V2 – DC sampling fan. CIXIKEJI DC brushless fan, model CX5010B5H, DC 5 V 0.23 A (label marks CE, FC, RoHS), with ferrule-terminated leads. |
| Figure 4-109. GLD V2 – Fan leads terminated. Fan leads connected to the FAN 5V/0V terminal. | Figure 4-110. GLD V2 – Alarm beacon fitted. Visual/audible alarm beacon (stainless-steel body, red dome) threaded into the right-hand entry; its leads are routed to the terminal board. |
| Figure 4-111. GLD V2 – Alarm leads terminated. Beacon leads connected to the ALARM terminal. | Figure 4-112. GLD V2 – Cable gland fitted to the left-hand entry. Cable gland of the current prototype (black polymer body) installed in the power-cable entry. |
| Figure 4-113. GLD V2 – 24 VDC cable inserted. 24 VDC supply cable passed through the cable gland. | Figure 4-114. GLD V2 – 24 VDC conductors terminated. Supply conductors connected to the 24V +/− terminal. |
| Figure 4-115. GLD V2 – Cover with sampling fan. Enclosure cover with the DC fan mounted directly behind the stainless-steel mesh; cover being fitted. | Figure 4-116. GLD V2 – Assembled unit — grounding. Cover closed; the external grounding screw beside the cable gland is being tightened. |
| Figure 4-117. GLD V2 – Threaded cover being closed. The threaded cover is screwed onto the base by hand. | Figure 4-118. GLD V2 – Assembled unit — front view. Stainless-steel mesh sensing face; cable gland (left), alarm beacon (right), antenna (top). |
| Figure 4-119. GLD V2 – Assembled unit — side view with blanking plug. Hexagon blanking plug closing the lower entry; alarm beacon on the right. | Figure 4-120. GLD V2 – Assembled unit — side view with cable entry. Cable gland, external grounding screw and antenna. |
| Figure 4-121. GLD V2 – Assembled unit — side view with beacon. Alarm beacon and mounting lugs. | Figure 4-122. GLD V2 – Assembled unit — side view of the base. Beacon entry (left), cable gland (right) and antenna. |
| Figure 4-123. GLD V2 – Assembled unit — rear view. Rear face of the base with mounting lugs; antenna (top), beacon (left), cable gland (right). |  |

#### 2.4.1 Photo-set coverage against the revision guide

The revision guide asks for a photo set covering twelve views. The table maps each requested view to the figures already included in this document and lists the views still to be photographed on the current production configuration (latest enclosure and PCB). Captions follow the format “Figure 4-n. GLD V2 – <title>” with a one-sentence description; no Ex marking is shown as final and the MQ sensor mesh is not described as flame-arresting.

| # | Requested view | Status | Where shown / action |
| --- | --- | --- | --- |
| 1 | Overall external view — front, rear, left, right, top, bottom | Available | Current production unit: Figures 4-118 (front), 4-119…4-122 (sides), 4-123 (rear), 4-116 (closed, top); CAD views M-A2 and M-B1…M-B4. |
| 2 | Product identification / nameplate area | Open | To be photographed once the nameplate position is fixed; caption to state that the marking is draft. |
| 3 | Front sensing area (mesh, cover, fan behind it) | Available | Figure 4-115 (fan behind the mesh in the cover), 4-118 (mesh face); M-D4. |
| 4 | Internal sensing section (8 MQ sensors, fan, chamber) | Available | Figures 4-107, 4-109…4-115 (sensor modules fitted, fan and cover); M-D6. |
| 5 | MQ sensor close-up (integrated protective mesh) | Available | Figure 4-105 (sensor module); M-D1. |
| 6 | Main PCB — component side and underside | Partial | Figures 4-102, 4-103 (component side in the enclosure); PCB layout views 2.6.a.8. Underside photograph to be added. |
| 7 | Power input and terminal area | Available | Figures 4-104 (labelled terminal board), 4-112…4-114 (gland, cable, 24V termination). |
| 8 | Grounding / bonding point | Partial | Figures 4-116, 4-120 (external grounding screw); close-up with lug and conductor to be added. |
| 9 | Cable entry and antenna interface (all penetrations) | Partial | Figures 4-112 (cable gland), 4-110 (beacon entry), 4-119 (blanking plug), 4-101 (antenna); close-up of the antenna bulkhead to be added. |
| 10 | Alarm module (LED/buzzer, gasket, fixing) | Available | Figures 4-110, 4-111; M-D5, M-E4. |
| 11 | Enclosure opened — complete internal arrangement | Available | Figure 4-115 (open base with sensors and terminals next to the cover with fan). |
| 12 | Exploded / assembly sequence photographs | Available | Figures 4-101…4-117 (complete assembly sequence); exploded illustrations M-D8, M-D9, M-E1. |

## 2.5 · Description of Intended Use and Installation Environment

The GLD is intended for continuous, fixed-point monitoring of flammable and toxic gas leaks in oil & gas refinery hazardous areas: process units, tank farms, pipe racks, and loading/unloading areas.

| Parameter | Value | Status |
| --- | --- | --- |
| Gas group | IIC | Internal recommendation, not an ExCB decision |
| Temperature class | T4 (≤135°C) | Target — preliminary internal measurement supports it (Section 2.6.f); formal verification outstanding |
| Area classification | Zone 1, Equipment Category 2G, Group II | Requested — subject to ExCB assessment |
| Type of protection | Ex d — flameproof enclosure | Internal recommendation, not an ExCB decision |

Type of protection selected over Ex e and Ex i based on the existing enclosure design: a die-cast aluminum (ADC12) housing with a stainless-steel mesh cover over the sensing element, functioning as a flame path/flame arrestor — the same pattern used by the three comparable certified gas detectors reviewed as references (all Ex d). Ex i was set aside because the MQ-series sensor heaters draw continuous power (≈0.75 W each) directly from an unregulated 24 VDC supply rather than through certified current-limiting barriers, difficult to reconcile with intrinsically-safe energy limits for Group IIC.

These classification parameters are the applicant's proposal, recorded here as a request, not as an agreed or granted classification. Cluster Head and Gateway devices are assumed to always sit in a safe area (project assumption, not the result of a formal area-classification study by Pertamina).

## 2.6 · Design and Manufacturing Information

### 2.6.a · Complete Drawings (Assembly, Component, Electrical Schematic, PCB Layout, Enclosure Structure, Junction Box, Terminal, Grounding)

An electrical schematic capture and a corresponding 9-sheet block-diagram set exist for the GLD V2 main board, exported from the native EasyEDA/JLCPCB source project, with pin-to-net traceability evidence.

Schematic block diagram, sheet 2 of 9 — main power distribution.

Main board PCB layout (top copper) — native EasyEDA/JLCPCB export, production-intent board.

A dimensioned mechanical drawing exists for the enclosure's external envelope and mounting hardware, drafted from a solid CAD model (STEP format, millimeter units):

Enclosure envelope & U-bolt mounting plate assembly — dimensioned drawing sheet (drafted 31 Aug 2026, from a STEP solid model).

A separate, internal contingency/alternate enclosure design (“GLD ATEX CASE v3”) exists in parallel with the primary commercial-enclosure sourcing path referenced in Section 2.3:

“GLD ATEX CASE v3” dimensioned drawing sheet (Farhan Budiman, 8 Sep 2026) — internal contingency/alternate enclosure design, not the confirmed production enclosure.

A candidate junction-box/terminal sub-component drawing has also been received via the casing partner's supply chain (Chinese enclosure-component manufacturer):

“Universal Base” BP18-1Z — supplier reference drawing (ADC12 die-cast aluminum alloy, 0.6 kg). Candidate component under evaluation; footprint does not match GLD's main enclosure, not yet confirmed as part of GLD's design.

| Drawing type | Status |
| --- | --- |
| Electrical schematic — main board | Available — 26 sheets PGLD-GLD-V4-MB-SCH-01 to -26, rev. 1.0 (Section 2.6.a.7); sensor-module board schematic not yet included |
| PCB layout — main board | Partially available — 10 views PGLD-GLD-V4-MB-PCB-01 to -10 (Section 2.6.a.8); layout rev/date, laminate data and sensor-module board layout open |
| Assembly drawing | Partially available — covers external envelope and mounting hardware, not internal PCB/component assembly sequence |
| Component drawing | Not yet available |
| Enclosure structure drawing (gap, length, volume) | Partially available — external envelope dimensioned; Ex-d-specific flame-path parameters not yet called out |
| Junction box / terminal / grounding connection drawings | Candidate reference only — BP18-1Z drawing above, not yet confirmed as part of GLD's design |

#### 2.6.a.1 Mechanical, enclosure and assembly drawings — complete set

This subsection reproduces every mechanical, enclosure and assembly drawing currently available for the GLD: the CAD drawings of the primary enclosure and mounting plate, the full sheet set of the GLD ATEX CASE v3 enclosure design, the supplier drawing of the candidate junction-box base, and the component, structure, terminal, grounding, assembly-sequence and sensor drawings of the manufacturer's design package. Each figure carries a reference number (M-xx) listed in the register below.

Drawing register

| Ref. | Title | Type | Group |
| --- | --- | --- | --- |
| M-A1 | ATEX CASING v2 — enclosure envelope & U-bolt mounting plate, dimensioned sheet | Dimensioned CAD drawing | A |
| M-A2 | ATEX CASING v2 — six orthographic CAD views (front, rear, left, right, top, bottom) | CAD render | A |
| M-B1 | GLD ATEX CASE v3 — sheet 1: isometric views of the assembled detector | CAD render | B |
| M-B2 | GLD ATEX CASE v3 — sheet 2: top and bottom views | CAD render | B |
| M-B3 | GLD ATEX CASE v3 — sheet 3: side and rear views | CAD render | B |
| M-B4 | GLD ATEX CASE v3 — sheet 4: front/back and mounting-lug views | CAD render | B |
| M-B5 | GLD ATEX CASE v3 — sheet 5: exploded views (four viewpoints) | CAD exploded view | B |
| M-B6 | GLD ATEX CASE v3 — sheet 6: exploded views of the sensor-head stack | CAD exploded view | B |
| M-B7 | GLD ATEX CASE v3 — sheet 7: dimensioned drawing with title block | Dimensioned CAD drawing | B |
| M-C1 | Universal Base BP18-1Z — supplier reference drawing | Supplier dimensioned drawing | C |
| M-D1 | LoRa antenna, MQ sensor and enclosure body & cover | Component photographs | D |
| M-D2 | Rubber gasket — locations | Component photographs | D |
| M-D3 | Rubber gasket — part | Component photograph | D |
| M-D4 | Stainless wire-mesh plate | Component photographs | D |
| M-D5 | Visual/audible alarm beacon; motherboard inside base casing | Component photographs | D |
| M-D6 | Motherboard with MQ sensors inside base casing | Component photograph | D |
| M-D7 | Common base design | Dimensioned drawing | D |
| M-D8 | Overall exploded views I and II | Exploded line drawing | D |
| M-D9 | Overall exploded view III | Exploded line drawing | D |
| M-D10 | Motherboard PCB power-in and PCB cover (terminal module) | CAD render | D |
| M-D11 | Motherboard and PCB cover | CAD render | D |
| M-D12 | Component drawing | Component photographs | D |
| M-D13 | Electrical block diagram and enclosure structure | Block diagram / dimensioned drawing | D |
| M-D14 | Terminal — 24 VDC power in | Photographs | D |
| M-D15 | Grounding | Photograph | D |
| M-E1 | Assembly drawing — exploded view | Assembly illustration | E |
| M-E2 | Assembly step — casing cover | Assembly illustration | E |
| M-E3 | Assembly step — sensor case and rubber gasket | Assembly illustration | E |
| M-E4 | Assembly step — LED & buzzer | Assembly illustration | E |
| M-E5 | Assembly step — 24 VDC power in | Assembly illustration | E |
| M-E6 | Assembly step — put motherboard into base casing | Assembly photographs | E |
| M-E7 | Assembly step — secure motherboard; install LoRa antenna cable | Assembly photographs | E |
| M-E8 | Assembly step — antenna position | Assembly photographs | E |
| M-E9 | Assembly step — alarm cables, DC fan and cover closure | Assembly photographs | E |
| M-F1 | MQ-2 and MQ-3 — sensor parameters and dimension drawings | Sensor data sheet | F |
| M-F2 | MQ-4 and MQ-5 — sensor parameters and dimension drawings | Sensor data sheet | F |
| M-F3 | MQ-6 and MQ-7 — sensor parameters and dimension drawings | Sensor data sheet | F |
| M-F4 | MQ-8 and MQ-135 — sensor parameters and dimension drawings | Sensor data sheet | F |

Status of this drawing set: the CAD sheets (M-A1, M-B7, M-C1, M-D7) carry real dimensions and title blocks; the remaining figures are CAD renders, line illustrations or photographs of production parts. Controlled, toleranced drawings that call out the Ex d flame-path parameters (joint gap and length, thread engagement, free internal volume) are still to be issued by the enclosure manufacturer. Two figures still show superseded details: M-E3 labels a battery case and M-D14 labels the power input as "battery or 24 VDC"; the production configuration is 24 VDC only. M-E8 records that the antenna has moved from the side to the top of the base enclosure.

A. Primary enclosure and mounting — ATEX CASING v2

M-A1. ATEX CASING v2 — enclosure envelope & U-bolt mounting plate, dimensioned sheet
Orthographic and isometric views with real dimensions (enclosure neck Ø75 mm, overall probe height 191.51 mm, mounting plate 250 × 250 mm with toleranced hole pattern), U-bolt parameter table (2″/DN50, M10 thread) and title block (drafted 31 Aug 2026). Drawn from the STEP solid model (mm).

| Front view | Rear view | Left view |
| --- | --- | --- |
| Right view | Top view | Bottom view |

M-A2. ATEX CASING v2 — six orthographic CAD views
Rendered from the same STEP solid model as M-A1 (enclosure, sensor head and U-bolt mounting plate).

B. Enclosure design GLD ATEX CASE v3 — complete sheet set

M-B1. GLD ATEX CASE v3 — sheet 1: isometric views of the assembled detector
Assembled enclosure with sensor head, external antenna and alarm beacon, three isometric viewpoints.

M-B2. GLD ATEX CASE v3 — sheet 2: top and bottom views
Top view (cable-entry bosses, alarm beacon, antenna) and bottom view (sensor-head mesh face).

M-B3. GLD ATEX CASE v3 — sheet 3: side and rear views
Left, right and rear-isometric views showing the base, sensor head, beacon and antenna positions.

M-B4. GLD ATEX CASE v3 — sheet 4: front/back and mounting-lug views
Views showing the wall-mounting lugs of the base and the cover opening.

M-B5. GLD ATEX CASE v3 — sheet 5: exploded views (four viewpoints)
Exploded assembly: base, cover, sensor case, mesh disc, sensor-case cover, beacon, antenna and glands.

M-B6. GLD ATEX CASE v3 — sheet 6: exploded views of the sensor-head stack
Sensor case, DC fan, stainless-steel filter mesh disc and case cover shown in stacking order.

M-B7. GLD ATEX CASE v3 — sheet 7: dimensioned drawing with title block
Dimensioned orthographic views and section of the sensor case; callouts for sensor case, DC fan, stainless-steel filter mesh disc, case cover and transparent glass; title block “GLD ATEX CASE v3” (drafted 8 Sep 2026). Internal contingency/alternate enclosure design.

C. Junction-box / terminal base — supplier drawing

M-C1. Universal Base BP18-1Z — supplier reference drawing
Die-cast base, material ADC12, 0.6 kg, scale 1:1: 2 × M20×1.5-6H cable-entry threads, 4 × M4-6H and 1 × M5×1.5-6g fixing threads, 3 × Ø8 mm mounting holes, internal bore up to Ø94.7 mm, overall ≈124 × 115 × 67.5 mm. Candidate junction-box/terminal sub-component from the casing supply chain.

D. Components, structure, terminal and grounding

M-D1. LoRa antenna, MQ sensor and enclosure body & cover
External LoRa antenna with cable, MQ sensor with integrated stainless-steel mesh, and the enclosure body and cover (production parts).

M-D2. Rubber gasket — locations
Rubber gaskets fitted at the cover-to-body and sensor-head joints (four locations marked).

M-D3. Rubber gasket — part
Rubber gasket as a separate part.

M-D4. Stainless wire-mesh plate
Front stainless-steel wire-mesh plate (gas inlet) and its seating in the cover.

M-D5. Visual/audible alarm beacon; motherboard inside base casing
LED and buzzer beacon with rubber gasket; main PCB mounted in the base enclosure.

M-D6. Motherboard with MQ sensors inside base casing
Main PCB with the eight MQ sensors fitted, seen through the open base enclosure.

M-D7. Common base design
Dimensioned base design drawing with title block.

M-D8. Overall exploded views I and II
Labelled exploded line drawings of the complete detector (cover, mesh, fan, sensor head, base, glands, beacon, antenna).

M-D9. Overall exploded view III
Exploded line drawing, third viewpoint.

M-D10. Motherboard PCB power-in and PCB cover (terminal module)
Populated main PCB (3D) and the PCB cover carrying the terminal module with the 24 VDC power-in position marked.

M-D11. Motherboard and PCB cover
Main PCB and PCB cover shown as separate parts.

M-D12. Component drawing
Enclosure; PCB + sensors; cover + fan; alarm module; cover + mesh; antenna 3 dBi.

M-D13. Electrical block diagram and enclosure structure
Block diagram of the main board and dimensioned enclosure-structure views.

M-D14. Terminal — 24 VDC power in
24 VDC waterproof cable connector and the terminal block on the PCB cover.

M-D15. Grounding
External grounding stud, marked with the protective-earth symbol, installed on the enclosure body.

E. Assembly sequence

M-E1. Assembly drawing — exploded view
Casing cover, stainless wire-mesh plate, wire-mesh locking bracket and DC fan.

M-E2. Assembly step — casing cover
Casing cover parts.

M-E3. Assembly step — sensor case and rubber gasket
Sensor case with rubber gasket (drawing also labels a battery case — see note).

M-E4. Assembly step — LED & buzzer
Alarm beacon fitted to the base.

M-E5. Assembly step — 24 VDC power in
Cable entry for the 24 VDC supply.

M-E6. Assembly step — put motherboard into base casing
Base casing and motherboard placement.

M-E7. Assembly step — secure motherboard; install LoRa antenna cable
Secure the motherboard with the designated mounting screws and install the antenna cable.

M-E8. Assembly step — antenna position
In the current version the LoRa antenna is positioned above the base enclosure (earlier side position superseded).

M-E9. Assembly step — alarm cables, DC fan and cover closure
Connect LED/buzzer cables; attach DC fan to the cover behind the stainless-steel mesh plate; connect the fan cable; close the cover by turning it clockwise.

F. MQ sensor dimension drawings

M-F1. MQ-2 and MQ-3 — sensor parameters and dimension drawings

M-F2. MQ-4 and MQ-5 — sensor parameters and dimension drawings

M-F3. MQ-6 and MQ-7 — sensor parameters and dimension drawings

M-F4. MQ-8 and MQ-135 — sensor parameters and dimension drawings

#### 2.6.a.7 Main PCB schematic

The electrical schematic of the GLD main board is provided as the actual circuit schematic (not a block diagram). It is presented as 26 figures, each cut from the original EasyEDA schematic around one functional block. Every component in a figure is enclosed in a numbered boundary box, and the same number is used in the description table printed with that figure. Each figure carries a sheet information table with the drawing/document number, revision, date, title and sheet number (n of 26).

Sheet register

| Drawing no. | Title | Requirement covered |
| --- | --- | --- |
| PGLD-GLD-V4-MB-SCH-01 | 24 VDC input, input protection and 24 V to 5 V buck (U36) | 24 VDC input; input protection |
| PGLD-GLD-V4-MB-SCH-02 | Power connector J1, battery input and external 5 V | 24 VDC input |
| PGLD-GLD-V4-MB-SCH-03 | Battery load switch (U14) | Input protection |
| PGLD-GLD-V4-MB-SCH-04 | 5 V power-path mux (U13) | 5 V rail |
| PGLD-GLD-V4-MB-SCH-05 | Battery boost to about 24 V (U1, TPS61175) | 5 V rail / +24 V backup |
| PGLD-GLD-V4-MB-SCH-06 | Battery boost to 5 V (U41, TPS61088) | 5 V rail |
| PGLD-GLD-V4-MB-SCH-07 | 5 V to 3.3 V buck (U43) | 3.3 V rail |
| PGLD-GLD-V4-MB-SCH-08 | Always-on 3.3 V LDO and its source selection (U42) | 3.3 V rail |
| PGLD-GLD-V4-MB-SCH-09 | TPL5010 timer/watchdog (U46) | 3.3 V rail |
| PGLD-GLD-V4-MB-SCH-10 | ENA latch (U51) and open-drain inverter (U15) | 3.3 V rail |
| PGLD-GLD-V4-MB-SCH-11 | 3V3AON voltage supervisor (U16) | 3.3 V rail |
| PGLD-GLD-V4-MB-SCH-12 | Sensor module connector H2 (one connector as example) | MQ heater supply |
| PGLD-GLD-V4-MB-SCH-13 | Sensor module connectors H1–H8 | MQ heater supply; sensor interface |
| PGLD-GLD-V4-MB-SCH-14 | Sensor module enable I/O expander (U12) | MQ heater supply (module enable) |
| PGLD-GLD-V4-MB-SCH-15 | VMID bias of about 2.5 V (U3) | MQ sensor analog circuitry |
| PGLD-GLD-V4-MB-SCH-16 | Analog +5VA supply filter (L6) | MQ sensor analog circuitry |
| PGLD-GLD-V4-MB-SCH-17 | Sensor module I2C multiplexer (U33) | MQ sensor analog circuitry |
| PGLD-GLD-V4-MB-SCH-18 | Temperature and humidity sensor (U44, SHT40) | Analog circuitry (environment) |
| PGLD-GLD-V4-MB-SCH-19 | ADS1256 ADC with input filters (AIN), clock and decoupling | ADC; analog circuitry |
| PGLD-GLD-V4-MB-SCH-20 | 2.5 V voltage reference and VREF buffer | ADC |
| PGLD-GLD-V4-MB-SCH-21 | ESP32-S3-WROOM-1U-N16R8 with support parts (incl. status LED) | ESP32-S3; LED |
| PGLD-GLD-V4-MB-SCH-22 | USB, CH340C and auto-reset | ESP32-S3 (programming) |
| PGLD-GLD-V4-MB-SCH-23 | E22-900MM22S LoRa module and antenna connector | LoRa |
| PGLD-GLD-V4-MB-SCH-24 | 5 V fan driver (Q5, J3) | Fan driver |
| PGLD-GLD-V4-MB-SCH-25 | 24 V alarm output driver (Q4, J2) | LED / buzzer (alarm) driver |
| PGLD-GLD-V4-MB-SCH-26 | RS-485 transceiver (U47) | RS-485 / Modbus |

All sheets: revision 1.0, date 2026-07-19.

Coverage of the required schematic content

| Required content (2.6.a.7) | Where shown | Status |
| --- | --- | --- |
| 24 VDC input | Sheets 01, 02 | Provided |
| Input protection | Sheets 01–03; protection-device table below | Provided — see findings 1–2 |
| 5 V rail | Sheets 01 (buck U36), 04–06 | Provided |
| 3.3 V rail | Sheets 07–11 | Provided |
| MQ heater supply | Sheets 12–14 (+5 V and enable to each sensor module) | Partly — heater switch is on the sensor module (finding 4) |
| MQ sensor bridge / analog circuitry | Sheets 15–18; AIN filters on sheet 19 | Partly — load resistor is on the sensor module (finding 4) |
| ADC | Sheets 19, 20 | Provided |
| ESP32-S3 | Sheets 21, 22 | Provided |
| LoRa | Sheet 23 | Provided |
| Fan driver | Sheet 24 | Provided |
| LED / buzzer driver | Sheet 25 (24 V alarm output J2); status LED on sheet 21 | Provided |
| RS-485 / Modbus | Sheet 26 | Provided |
| Protection devices | Table below | Provided |
| Grounding | Text below | Provided — PE point not drawn (finding 5) |
| Title block per sheet | Sheet information table on every figure | Provided |

Protection devices

| Designator | Part | Location | Function | Sheet |
| --- | --- | --- | --- | --- |
| F1 | MINISMDC260F/16-2 | Series with 24V+ after J1 | Overcurrent on 24 V line (PTC) | 01 |
| D1 | SMBJ33A | After F1, to 24V− | Surge / transient clamp on 24 V line (TVS) | 01 |
| L1 | ACM7060-301-2PL | Series with 24V+ and return | Common-mode noise | 01 |
| L4 | MPZ2012S101AT000 | Series with +24V | High-frequency noise (ferrite) | 01 |
| D2 / R26 | LBZT52C12T1G / 100k | Gate of Q3 | 12 V gate-source limit for Q3 | 01 |
| Q3 | SI7465DP-T1-GE3 | +24V to buck input | High-side switch | 01 |
| F2 | MINISMDC260F/16-2 | Series with battery input J1.2 | Battery overcurrent (PTC) | 02 |
| Q2 / R1 | AO4407 / 100k | Battery to VBAT_IN | High-side switch | 02 |
| D7 / D13 | SS54 | 5VEXT to +5V; boost to +24V | Reverse-current blocking (OR-ing) | 02, 05 |
| D3 / D10 | BAT54C | Source of U42 (3V3AON) | OR-ing of supply sources | 08 |
| D8 / D9 | SS14 | Loads J2 and J3 | Flyback for inductive loads | 24, 25 |
| D4 / D5 / D11 | LESD5D5.0CT1G | USB D−, D+, VBUS | USB ESD | 22 |
| D6 | SM712 | RS-485 A/B | RS-485 line TVS | 26 |
| L6 | GZ1005D600TF | +5V to +5VA | Analog supply noise isolation | 16 |

Grounding

The circuit uses a single GND net; there is no separate AGND or chassis-ground net. The AGND/PGND/DGND pins of all ICs are connected to GND.

The 24 V return (J1 pin 5, 24V−) enters GND through one winding of common-mode choke L1 (pin 2 to pin 3). J1 pins 1 and 3 are connected directly to GND.

The analog supply +5VA is separated from +5V by ferrite bead L6; analog and digital ground are joined at GND.

The shell/EP pads of USB1 and the GND pads of the LoRa module/antenna connector are connected to GND.

A connection point from circuit GND to the enclosure / protective earth is not drawn on the schematic; it is defined by the enclosure grounding/bonding drawing.

Findings to be resolved before the dossier is finalised

| # | Finding (from the schematic) | Action |
| --- | --- | --- |
| 1 | F1 and F2 are MINISMDC260F/16-2; the "/16" suffix normally means a 16 V rating, while F1 is on the 24 V line. | Confirm the PTC voltage rating from the datasheet or change the part. |
| 2 | Q3 (24 V) and Q2 (battery) have the source on the input side and the gate pulled to GND; with this orientation the body diode does not block reverse polarity. | Confirm the intended function. Any description of "P-channel MOSFET reverse-polarity protection" must match the actual circuit. |
| 3 | Battery input, battery boost to 24 V (TPS61175) and to 5 V (TPS61088), battery load switch and power-path mux are populated on the main board. | Decide whether the battery path is excluded (depopulated / not fitted) or included in the certified configuration; the production configuration is stated as 24 VDC only. |
| 4 | Each sensor connector H1–H8 carries only +5 V, VMID, enable, I2C and AIN. The per-channel heater switch and the sensor load resistor are not on the main board. | Provide the sensor-module schematic and layout as a separate drawing set. |
| 5 | Single GND net; no chassis/PE connection point on the board. | Show the GND-to-enclosure/PE arrangement in the grounding/bonding drawing. |
| 6 | RS-485 has TVS D6 only; no 120 Ω termination or A/B bias resistors. | Confirm termination/bias is provided externally (installation note) or is not required. |
| 7 | Converter output voltages (≈5.0 V, ≈24 V, ≈5.3 V) are calculated from feedback resistors, not measured. | Record measured rail voltages in the routine/functional test. |
| 8 | The load type on alarm output J2 (24 V siren/beacon) is not defined on the schematic. | Identify the alarm module (manufacturer/model, current) in the BOM (6.b). |

Schematic figures (sheets 01–26)

The pages below reproduce the schematic figures exactly as issued, including the per-figure sheet information table and component description table.

#### 2.6.a.8 PCB layout

The layout of the GLD main board is presented in ten views generated from the EasyEDA PCB file of the same project. Each view highlights one aspect required by item 2.6.a.8; numbered markers on the view correspond to the legend printed below it.

Board data

| Parameter | Value | Source / status |
| --- | --- | --- |
| Board name | Main Board GLD (MotherBoardGLDVer2) | Layout file |
| Board outline | Circular, Ø 84.0 mm | Dimension in view 04 |
| Copper layers | Top and bottom (two views of copper) | Views 01, 02 |
| Component side | All SMD parts on top; only through-hole connectors pass to the bottom | View 02 |
| Enclosure fixing holes | 4 × Ø 3.2 mm on Ø 64 mm pitch circle, at (0, +32), (+32, 0), (0, −32), (−32, 0) mm from centre | View 05 |
| Additional holes | 4 × Ø 3.0 mm near the board edge (≈ ±39 mm from centre) | View 05 — purpose (casing screw or standoff) to be confirmed against the mechanical drawing |
| Ground | GND pour on top and bottom; 176 of 400 vias stitch the GND pours | Views 09, 10 |
| Laminate, thickness, copper weight | Not stated in the layout package | To be provided (FR-4 grade, UL 94 rating, Tg, CTI) — see 6.b / 6.c |
| Layout revision / date | Not stated in the layout package | To be added to the drawing table |

Drawing register

| Drawing no. | Title | Requirement covered | Rev / date |
| --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-01 | Top layer | Top layer | TBC |
| PGLD-GLD-V4-MB-PCB-02 | Bottom layer | Bottom layer | TBC |
| PGLD-GLD-V4-MB-PCB-03 | Component placement | Component placement | TBC |
| PGLD-GLD-V4-MB-PCB-04 | Board dimensions | Board dimensions | TBC |
| PGLD-GLD-V4-MB-PCB-05 | Mounting holes | Mounting holes | TBC |
| PGLD-GLD-V4-MB-PCB-06 | Connectors | Connectors | TBC |
| PGLD-GLD-V4-MB-PCB-07 | High-current / power area | High-current / power area | TBC |
| PGLD-GLD-V4-MB-PCB-08 | Sensor interface | Sensor interface | TBC |
| PGLD-GLD-V4-MB-PCB-09 | Grounding arrangement — top | Grounding arrangement | TBC |
| PGLD-GLD-V4-MB-PCB-10 | Grounding arrangement — bottom | Grounding arrangement | TBC |

Coverage of the required layout content

| Required content (2.6.a.8) | Where shown | Status |
| --- | --- | --- |
| Top layer | View 01 | Provided |
| Bottom layer | View 02 | Provided |
| Component placement | View 03 | Provided |
| Board dimensions | View 04 | Provided |
| Mounting holes | View 05 | Provided — purpose of 4 × Ø 3.0 mm holes to confirm |
| Connectors | View 06 | Provided |
| High-current / power area | View 07 | Provided |
| Sensor interface | View 08 | Provided |
| Grounding arrangement | Views 09, 10 | Provided |
| One set per board (Main, Sensor, Alarm, Terminal) | Main board only | Open — sensor-module board layout to be added |

Figure 8-1. Top layer — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-01 | TBC | TBC | Top layer | 1 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | ESP32-S3 | Main MCU module: runs the GLD firmware; SPI to LoRa and ADC, I²C to sensors, GPIO for alarm/fan. |
| 2 | LoRa E22-900MM22S | LoRa radio for communication with the cluster head. |
| 3 | U.FL antenna connector | Connection to the external LoRa antenna. |
| 4 | Sensor ports (8) | Eight 2 × 4 headers around the board, one per sensor module (yellow circle = module outline). |
| 5 | ADS1256 | 24-bit ADC reading the eight analog sensor channels. |
| 6 | TCA9548A | 8-channel I²C multiplexer, one channel per sensor port. |
| 7 | PCF8574 | I/O expander providing the enable signals EN0–EN7 for each sensor. |
| 8 | Micro-USB | USB port for firmware flashing and serial debug. |
| 9 | CH340C | USB-to-UART converter for programming and serial monitor. |
| 10 | ALARM connector | 24 V siren/beacon output. |
| 11 | FAN connector | 5 V fan output. |
| 12 | RS-485 connector | RS-485 bus (A/B). |
| 13 | VIN connector | Power input: 24V+, 24V−, 5VEXT, GND, BAT (battery). |
| 14 | PTC fuse | Limits current on the 24 V line under short circuit. |
| 15 | TVS | Clamps voltage surges on the 24 V input. |
| 16 | Common-mode filter | Filters common-mode noise on the 24 V input. |
| 17 | Buck 24 V → 5 V | LMR51450 regulator stepping 24 V down to 5 V. |
| 18 | Battery boost → 24 V | TPS61175 boosting the battery to the +24V rail. |
| 19 | Battery boost → 5 V | TPS61088 generating 5 V from the battery. |

Figure 8-2. Bottom layer — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-02 | TBC | TBC | Bottom layer | 2 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | Bottom GND pour (solid blue area) | GND plane filling almost the whole bottom side; connected to the top pour through vias. |
| 2 | Through-hole connectors | The only parts passing through to the bottom; all SMD parts are on the top side. |

Figure 8-3. Component placement — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-03 | TBC | TBC | Component placement | 3 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | ESP32-S3 | Main MCU: runs the firmware and controls all blocks. |
| 2 | LoRa (E22-900MM22S) | LoRa radio for communication with the cluster head. |
| 3 | Antenna (U.FL) | External LoRa antenna connector. |
| 4 | ADS1256 | 24-bit ADC for the eight analog sensor channels. |
| 5 | TCA9548 | 8-channel I²C mux, one channel per sensor port. |
| 6 | PCF8574 | I/O expander for the enable signals EN0–EN7. |
| 7 | USB | Micro-USB + CH340C for firmware flashing and serial monitor. |
| 8 | RS-485 | THVD1410 transceiver for the RS-485 bus. |
| 9 | Alarm driver | MOSFET switching the 24 V siren/beacon. |
| 10 | Fan driver | MOSFET switching the 5 V fan. |
| 11 | SHT40 | Temperature and humidity sensor inside the enclosure. |
| 12 | Power block | Power input and protection (VIN connector, fuse, TVS, filter). |
| 13 | TPL5010 | Watchdog timer that wakes the system periodically. |

Figure 8-4. Board dimensions — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-04 | TBC | TBC | Board dimensions | 4 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | Board diameter: 84 mm | Circular board (outline) matching the round aluminium enclosure. |
| 2 | Enclosure-hole pitch circle: Ø 64 mm | The four enclosure fixing holes lie on this circle, 90° apart. |

Figure 8-5. Mounting holes — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-05 | TBC | TBC | Mounting holes | 5 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | Enclosure holes (4 × Ø 3.2 mm) | Hole pattern of the enclosure footprint on the Ø 64 mm circle; board-to-enclosure fixing screws. Positions from centre: (0, +32), (+32, 0), (0, −32), (−32, 0) mm. |
| 2 | Additional holes (4 × Ø 3.0 mm) | Four further fixing holes near the board edge (≈ ±39 mm from centre). Exact use (enclosure screw or standoff) to be confirmed with the mechanical drawing. |

Figure 8-6. Connectors — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-06 | TBC | TBC | Connectors | 6 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | VIN | Power input, 2 × 3 pins: 24V+, 24V−, 5VEXT, GND, BAT (battery). |
| 2 | ALARM | 24 V siren/beacon output (pin 1 = +24V, pin 2 = MOSFET switch side). |
| 3 | FAN | 5 V fan output (pin 1 = +5V, pin 2 = MOSFET switch side). |
| 4 | RS-485 | RS-485 bus: B (pin 1) and A (pin 2). |
| 5 | Micro-USB | Firmware flashing and serial debug. |
| 6 | U.FL (antenna) | External LoRa antenna connector. |
| 7 | Sensor ports (8) | Eight 2 × 4 headers (1.27 mm pitch): GND, +5V, AINx, SCLx, SDAx, ENx, VMID — one sensor module each. |

Figure 8-7. High-current / power area — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-07 | TBC | TBC | High-current / power area | 7 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | Power input & protection | VIN connector, fuse, TVS, common-mode filter, ferrite: 24 V entry path. |
| 2 | 24 V → 5 V converter | LMR51450 buck, 4.7 µH inductor, P-MOSFET switch. |
| 3 | Battery → 24 V converter | TPS61175 boost, 10 µH inductor, SK36 diode and OR-ing diode to the +24V rail. |
| 4 | Battery → 5 V converter + power mux | TPS61088 boost, 2.2 µH inductor, TPS2116 power mux selecting the 5 V source. |
| 5 | Battery protection | Fuse, P-MOSFET, load switch: disconnects the battery when the system sleeps. |

Figure 8-8. Sensor interface — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-08 | TBC | TBC | Sensor interface | 8 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | Sensor ports (8) | Eight headers: +5V, GND, VMID, ENx, SDAx/SCLx (I²C per channel), AINx (analog to ADC). |
| 2 | TCA9548A | I²C mux splitting the ESP32 I²C bus into 8 channels for 8 sensors. |
| 3 | PCF8574 | I/O expander generating EN0–EN7 to switch each sensor module. |
| 4 | ADS1256 + 8 MHz crystal | 8-channel 24-bit ADC, SPI to ESP32, clocked from an 8 MHz crystal. |
| 5 | 2.5 V reference + buffer | ADR03 generates VREF, buffered by OPA320 for the ADC. |
| 6 | VMID buffer | OPA320 follower generating VMID (analog mid-bias) for the sensor modules. |

Figure 8-9. Grounding arrangement (top) — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-09 | TBC | TBC | Grounding arrangement — top | 9 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | GND pour | One GND copper area (green line along the board edge) on top and bottom, filling all free area: short return paths and noise shielding. |
| 2 | GND vias (examples) | 176 of the 400 vias connect the top and bottom GND pours; blue circles mark examples. |
| 3 | 24V− → GND at the common-mode filter | The 24 V input return (24V−) joins system GND through the return winding of the common-mode choke. |
| 4 | Analog area without split ground | Single GND net. The analog area (ADS1256, voltage reference, OPA320) is separated through the +5VA supply (ferrite), not by a split ground. |

Figure 8-10. Grounding arrangement (bottom) — Main Board GLD

| Drawing no. | Rev | Date | Title | Sheet |
| --- | --- | --- | --- | --- |
| PGLD-GLD-V4-MB-PCB-10 | TBC | TBC | Grounding arrangement — bottom | 10 of 10 |

| No | Item | Description |
| --- | --- | --- |
| 1 | GND pour | Bottom GND copper area filling all free area. |
| 2 | GND vias (examples) | 176 of 400 vias stitch the top and bottom GND pours; blue circles mark examples. |

Open items for 2.6.a.8: (a) layout revision and date for the drawing table; (b) PCB laminate grade, UL 94 rating, Tg, CTI, board thickness and copper weight; (c) purpose of the 4 × Ø 3.0 mm holes; (d) layout drawings of the sensor-module board.

### 2.6.b · Bill of Materials (BOM) for Key Components Affecting Explosion-Proof Safety

A complete, itemized electronic-component BOM exists for both boards, exported directly from the EasyEDA/JLCPCB source project (manufacturer, manufacturer part number, and LCSC supplier part number per line item): 89 line items / 214 placed components on the main board, 14 line items / 24 placed components on the sensor board. Full listing maintained in Dokumen_Teknis_Sertifikasi_GLD_IECEx_ATEX, Section 2.6.b.

Status: Partially available. The explosion-safety-relevant subset the checklist actually asks for (enclosure, gasket, cable entry device, battery, potting compound, and the gas sensor itself, with material grade and Ex/UL/CCC certification for each) remains not yet compiled — these are mechanical/safety parts, not electronic components, and sit with the Manufacturer, PT Galaksi Megatama Indonesia. The gas sensor itself (MQ2) is sourced outside the LCSC supply chain, without a manufacturer/LCSC part reference; its Ex status is not yet verified.

#### 2.6.b.1 Ex-critical mechanical / construction components

Components whose substitution could affect the explosion-protection concept, in the format requested by the revision guide. The electronic BOM above remains the full component list; this table is the Ex-critical subset and is controlled by part number or drawing number.

| Ref. | Component | Manufacturer | Part / drawing no. | Material / rating | Safety function | Evidence | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| EX-01 | Main enclosure body | PT Galaksi Megatama Indonesia | GLD ATEX CASE v3 / common base (M-B7, M-D7) | ADC12 die-cast aluminium alloy; wall thickness TBC; surface treatment TBC | Flameproof enclosure (Ex d, proposed) | Drawings M-B7, M-D7 | Open |
| EX-02 | Enclosure cover (threaded) | PT Galaksi Megatama Indonesia | TBC | Aluminium alloy; thread form, pitch and engagement TBC | Flameproof joint | Assembly M-E9 (closed by turning clockwise) | Open |
| EX-03 | Front stainless-steel wire-mesh plate + locking bracket | TBC | TBC (M-D4, M-E1) | Stainless steel; grade, mesh/pore size, thickness, layers TBC | Gas inlet; flame-path element if certified | Photograph, assembly illustration | Open |
| EX-04 | MQ sensor integrated protective mesh | Sensor manufacturer TBC | MQ-2, MQ-3B, MQ-4, MQ-5, MQ-6, MQ-7B, MQ-8, MQ-135 | Stainless-steel mesh (per sensor construction) | Protective element; no flame-arresting claim | Sensor data sheets (M-F1…M-F4) | Open |
| EX-05 | DC sampling fan | CIXIKEJI | CX5010B5H (DC brushless, 50 × 50 × 10 mm frame) | DC 5 V, 0.23 A (≈1.2 W); CE/FCC/RoHS only — no Ex certification; max. temperature TBC | Gas sampling; potential ignition source (IS-02) | Label, Figure 4-108; driven by Q5 (sheet 24) | Partial |
| EX-06 | Gaskets / O-rings | TBC | TBC (M-D2, M-D3) | Rubber; compound, hardness, temperature range TBC | Environmental sealing (IP66) | Photographs | Open |
| EX-07 | Cable gland (24 VDC entry) | TBC | TBC | M20×1.5 entry thread (BP18-1Z reference); Ex d barrier gland required for IIC; cable range, IP, temperature TBC | Cable-entry flame path and sealing | Current prototype uses a general-purpose black polymer gland (Figure 4-112) — to be replaced by a certified metallic Ex d barrier gland | Pending certificate |
| EX-08 | Blanking plug / adaptor | TBC | TBC | Metallic hexagon plug fitted in the lower entry; thread and certification TBC | Closes unused entry | Figure 4-119 | Pending certificate |
| EX-09 | Antenna bulkhead / SMA interface | TBC | TBC | SMA; material and sealing TBC | Enclosure penetration | — | Open |
| EX-10 | External antenna | TBC | TBC | Omnidirectional, 3 dBi, SMA; material/environmental rating TBC | RF radiator | M-D1 | Open |
| EX-11 | LED / buzzer alarm beacon | TBC | TBC | Driven at 24 V from J2 (sheet 25); housing, gasket, current TBC | Local alarm; separate penetration | M-D5 | Open |
| EX-12 | Grounding stud | TBC | TBC | External stud with PE symbol; size and material TBC | Protective earth / bonding | M-D15; Figures 4-116, 4-120 | Open |
| EX-13 | Mounting fasteners | TBC | U-bolt mounting plate (M-A1) | U-bolt 2″/DN50, M10 thread; plate 250 × 250 mm; fastener grade TBC | Mechanical retention | Drawing M-A1 | Partial |
| EX-14 | Internal insulating parts (PCB cover / terminal module) | TBC | M-D10, M-D11 | Material, flammability, CTI TBC | Insulation of terminals | CAD render | Open |
| EX-15 | PCB substrate | TBC (PCB fabricator) | Main board Ø84 mm (2.6.a.8) | Laminate grade, UL 94 rating, Tg, CTI, thickness TBC | Carrier of energized circuits | Layout views | Open |
| EX-16 | Potting / adhesive | — | — | Not used in the current design (to be confirmed) | — | — | TBC |
| EX-17 | Battery and holder | — | Battery path on main board (sheets 02–06); BAT −/+ terminal on the terminal board (Figure 4-104); battery case in M-E3 | — | Included only if retained in the certified configuration | Schematic | Decision required |

2.6.b.2 Ex-critical electrical components

Electrical components that affect input protection, fault energy, heater power, surface temperature, isolation, switching or ignition risk. Data from the released main-board schematic (2.6.a.7) and the electronic BOM.

| Ref. | Component | Manufacturer | Part no. | Rating | Safety function | Evidence | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| EL-01 | Resettable PTC fuse F1 (24 V) / F2 (battery) | Littelfuse | MINISMDC260F/16-2 | I-hold 2.6 A, I-trip 5.0 A, V-max 16 V | Overcurrent protection | Datasheet; sheets 01, 02 | Review — 16 V rating on a 24 V line |
| EL-02 | TVS diode D1 | Ruilon | SMBJ33A | V-RWM 33 V, P-pk 600 W | Input transient/overvoltage | Datasheet; sheet 01 | Final |
| EL-02a | TVS D6 (RS-485), ESD D4/D5/D11 (USB) | DOWO / UMW | SM712 / LESD5D5.0CT1G | Line clamps | Interface protection | Sheets 22, 26 | Final |
| EL-03 | P-MOSFET Q3 + zener D2 | Vishay / LRC | SI7465DP-T1-GE3 / LBZT52C12T1G | V-DS, I-D per datasheet; V-GS limited to 12 V | Input switch (reverse-polarity function TBC) | Sheet 01 | Review — orientation |
| EL-03a | Common-mode choke L1, ferrite L4 | TDK | ACM7060-301-2PL / MPZ2012S101AT000 | Per datasheet | Input EMI filtering | Sheet 01 | Final |
| EL-04 | Buck converter U36 (24 V → 5 V) | Texas Instruments | LMR51450 | VIN ≤ 36 V; output ≈5.0 V (calculated) | Power conversion; possible hot surface | Sheet 01 | Final |
| EL-04a | Buck converter U43 (5 V → 3.3 V), LDO U42 | Texas Instruments | TPS62162 / TPS7A0233 | 3.3 V rails | Logic supply | Sheets 07, 08 | Final |
| EL-05 | Boost converters U1 (batt → 24 V), U41 (batt → 5 V) | Texas Instruments | TPS61175 / TPS61088 | Battery path | Power conversion; possible hot surface | Sheets 05, 06 | Decision required — battery path |
| EL-05a | Battery load switch U14, power mux U13, P-MOSFET Q2 | TI / JSMSEMI | TPS22964C / TPS2116 / AO4407 | Battery path | Battery switching | Sheets 02–04 | Decision required — battery path |
| EL-06 | Power inductors | ABC / SXN / APV | GSSM06304R7M2AU 4.7 µH; SMDRH105R-2R2NT; ANR3015T2R2M; SMDRH105R-100MT | I-sat, I-rms, ΔT TBC | Possible hot surface | Sheets 01, 05, 06, 07 | Open — ΔT not measured |
| EL-07 | MQ heater supply | — | +5 V to sensor connectors H1–H8; enable via PCF8574 (U12) | Heater voltage 5 V; per-channel switch on sensor module | Heater power / surface temperature | Sheets 12–14 | Open — sensor-module schematic |
| EL-08 | Fan driver Q5 + flyback D9 | — / MDD | AO3400A / SS14 | Low-side switch, 5 V | Fan current | Sheet 24 | Final |
| EL-09 | Alarm driver Q4 + flyback D8 | — / MDD | AO3400A / SS14 | Low-side switch, 24 V output J2 | Alarm load switching | Sheet 25 | Final |
| EL-10 | Main PCB | PCB fabricator TBC | MotherBoardGLDVer2 | Ø84 mm, two copper layers | Carrier of energized circuits | 2.6.a.8 | Open — laminate data |
| EL-11 | ESP32-S3 module | Espressif | ESP32-S3-WROOM-1U-N16R8 | 3.3 V | Main energized electronics | Sheet 21 | Final |
| EL-12 | LoRa module | Ebyte | E22-900MM22S | 0–22 dBm, 920–923 MHz | RF source | Sheet 23 | Final |
| EL-13 | Terminals / connectors | Boomele / TBC | J1 power 2×3 (pin 6 = 24V+, pin 5 = 24V−); XH-2A; H1–H8 2×4 1.27 mm | Current/voltage rating TBC | Loose-contact heating / sparking | Sheets 02, 13 | Open — VIN connector MPN |

Status legend: Final = confirmed and documented; Review = data exists but a rating or function must be checked; Open = data not yet available; TBC = to be confirmed; Pending certificate = certificate to be obtained; Decision required = design decision needed; N.A. = not applicable.

### 2.6.c · Material Specification Sheets / Datasheets (Non-Metallic Materials)

Datasheets or supplier conformity declarations for non-metallic materials (enclosure components, seals, insulators, potting compounds) — covering heat/cold resistance, anti-aging, anti-static, flame retardancy, CTI value, and chemical resistance — have not yet been collected.

Per IEC 60079-0:2017, once GLD's specific non-metallic components are identified the applicable requirements are: a Comparative Tracking Index (CTI) class per Table 5/6 appropriate to the equipment's creepage/clearance design; a glow-wire or needle-flame flame-retardancy result per §26.4–26.5; a glass-transition temperature (Tg) safely outside the confirmed −20°C to +60°C ambient operating range (Section 2.3); and UV/weathering-resistance data for any component exposed outdoors. These thresholds are cited from the standard itself and have not yet been matched against actual component datasheets.

Non-metallic components in the design that this item will need to cover: the cover-to-body sealing gasket providing the IP66 rating (Section 2.3 — identified as a rubber gasket in the casing partner's assembly drawing, specific compound/grade not yet specified); the integral seal of the candidate M20×1.5 cable gland — per project direction, specified together with the gland itself from the BP18-1Z reference (Section 2.3), rather than as a separately sourced component; and the antenna/SMA feedthrough insulator (material not yet specified). The protective mesh cover over the gas-sensing element is now identified as a stainless-steel wire mesh in the same assembly drawing — being metallic, it falls outside this specific non-metallic-materials checklist item.

Status: Not yet available, pending the enclosure design and material selection with the casing partner — this data has been formally requested from that partner. Partial candidate data: the BP18-1Z reference drawing (Section 2.6.a) specifies a powder-coating surface treatment with electrostatic-safety parameters directly relevant to this item — maximum surface charge transfer <10 nC and maximum surface capacitance <5 pF (coating thickness ≤0.2 mm), consistent with IEC 60079-0 coating requirements — but this describes a supplier's candidate component, not a confirmed datasheet for GLD's own enclosure finish.

#### 2.6.c.1 Non-metallic materials register (MAT-01 … MAT-11)

Format requested by the revision guide. Each material is to be supported by the supplier's datasheet or conformity declaration covering service-temperature range, heat/cold resistance, ageing, flame retardancy (UL 94 / glow-wire), CTI where relevant, electrostatic properties, chemical and UV/weather resistance, and Tg where relevant. Where the material is not yet identified the item stays Open.

| Ref. | Component | Material / grade | Manufacturer | Important properties | Evidence | Status |
| --- | --- | --- | --- | --- | --- | --- |
| MAT-01 | Cover-to-body gasket | Rubber; compound TBC | TBC | Temp. range, ageing, chemical resistance | Datasheet | Open |
| MAT-02 | Cable-gland body and sealing ring | Current prototype: black polymer gland (Figure 4-112); final: per certified Ex d gland | Gland supplier | Temp., flame rating, UV/chemical, electrostatic | Datasheet / certificate | Open |
| MAT-03 | Main PCB laminate | FR-4 type expected; exact grade TBC | PCB fabricator | Tg, CTI, UL 94 | Material datasheet | Open |
| MAT-04 | Connector housings (J1, XH-2A, H1–H8) | TBC | Boomele / TBC | CTI, UL 94, temp. rating | Datasheet | Open |
| MAT-05 | Wire insulation (24 VDC cable, 2 conductors L+/L−, Ø≈0.75 mm) | TBC | TBC | Temp. rating, flame rating | Cable datasheet | Open |
| MAT-06 | Antenna housing / insulator | TBC | TBC | Temp., UV, electrostatic properties | Datasheet | Open |
| MAT-07 | Fan impeller / housing | Plastic (grade TBC) | CIXIKEJI (CX5010B5H) | Temp., flame rating, mechanical | Fan datasheet | Open |
| MAT-08 | Adhesive / sealant | Not used (TBC) | — | Temp., chemical resistance, ageing | Datasheet | TBC |
| MAT-09 | Enclosure powder coating | Powder coating ≤0.2 mm (BP18-1Z reference) | Casing supply chain | Charge transfer <10 nC; surface capacitance <5 pF | Supplier drawing | Partial |
| MAT-10 | Transparent window (GLD ATEX CASE v3) | Glass; type TBC | TBC | Impact and thermal-shock tests for Ex d windows; temp. rating | Drawing M-B7 | Open |
| MAT-11 | Alarm beacon lens / dome | TBC | TBC | UV, impact, electrostatic, flame rating | Datasheet | Open |

### 2.6.d · Manufacturing Process Description

A description of manufacturing processes relevant to explosion-protection safety (enclosure machining accuracy control, explosion-proof surface treatment, welding, potting, die-casting, bonding) has not yet been documented.

Status: Not yet available — this describes the Manufacturer's (PT Galaksi Megatama Indonesia) process, not an internal electronics process.

#### 2.6.d.1 Manufacturing and assembly process control

Manufacturing and assembly processes that may affect explosion protection are controlled through approved drawings, work instructions, inspection criteria and production records. The sub-sections below follow the structure of the revision guide; parameters that the enclosure manufacturer has not yet issued are marked TBC rather than assumed.

d.1 Enclosure manufacturing and machining control

Body and cover are produced by die casting (ADC12) and machined to controlled drawings. Safety-critical dimensions — threaded joints, wall thickness, sealing surfaces and other flame-path features — are inspected against the specified tolerances before assembly. Inspection method, measuring equipment and acceptance criteria: TBC (enclosure manufacturer). For Ex d, gap, joint length and thread engagement are the controlled characteristics.

d.2 Enclosure surface treatment

Powder coating is shown on the BP18-1Z reference (≤0.2 mm, charge transfer <10 nC, surface capacitance <5 pF). Coating of flame-path surfaces is not permitted unless approved. Final treatment for the GLD enclosure: TBC.

d.3 Front stainless-steel mesh assembly

The mesh plate is fitted with its locking bracket in the specified orientation (M-E1) and inspected for deformation, damage, contamination or incomplete seating before the cover is accepted. Mesh part number, fastening torque and any adhesive: TBC.

d.4 DC fan installation

The DC fan is attached to the enclosure cover behind the stainless-steel mesh plate and its cable connected to J3 (M-E9). Airflow direction, screw type, torque, clearance to the mesh and connector locking are to be defined on the assembly drawing (TBC).

d.5 MQ sensor array installation

Eight MQ sensors (MQ-2, MQ-3B, MQ-4, MQ-5, MQ-6, MQ-7B, MQ-8, MQ-135) are mounted on the sensor modules at fixed positions H1–H8. Position/type correspondence is verified visually to prevent interchange; the integrated sensor mesh must be undamaged. Position map: TBC on a controlled drawing.

d.6 PCB assembly and inspection

SMT assembly by the PCB assembler from the released design files (EasyEDA/JLCPCB), followed by visual inspection, polarity/orientation check, firmware programming and functional test. PCB revision is recorded for traceability. Assembler name/address and AOI use: TBC.

d.7 Cable entry and terminal assembly

Only the approved gland/adaptor for the M20×1.5 entry is used; thread verification, sealing ring, cable size, tightening torque and treatment of unused entries (approved blanking plug) are checked. The 24 VDC supply is connected to J1 pins 6 (24V+) and 5 (24V−). Torque values: TBC.

d.8 Grounding and bonding assembly

External grounding stud with washer–lug–nut sequence; continuity between stud and all conductive enclosure parts is tested before release. Stud size, torque and continuity acceptance limit: TBC.

d.9 Gasket and sealing installation

Correct gasket part, orientation, clean and undamaged sealing surfaces, no twisting; replacement whenever damaged or after a defined number of openings (TBC).

d.10 Final enclosure closure

Clean mating/threaded surfaces, close the cover by turning clockwise to full engagement, apply the locking feature and inspect after closure. Minimum thread engagement and final torque: TBC.

d.11 Firmware and configuration control

Each unit is programmed with the approved firmware version, AI model version (three classes: clean air, LPG, H2), configuration file, alarm thresholds and LoRa settings. Programming is verified and recorded against the serial number. Changes require controlled approval and a rollback path.

d.12 Final inspection and routine test

Visual inspection; enclosure integrity; mesh condition; fan operation; 8/8 sensor detection; power consumption at 24 VDC (≤8 W); alarm test; LoRa communication; grounding continuity; IP/leak check if required as a routine test; serial number and documentation check. Ex d routine overpressure test, if required by the certificate: TBC.

d.13 Applicability of special processes

| Process | Applicability |
| --- | --- |
| Die casting | Applicable — enclosure in ADC12 die-cast aluminium |
| Machining of flame-path surfaces | Applicable for Ex d (final concept to be confirmed) |
| Welding | Not used in the current design (to be confirmed by the enclosure manufacturer) |
| Potting | Not used (to be confirmed) |
| Bonding adhesive | Not used unless specified for mesh or seal (TBC) |
| Surface treatment (powder coating) | Applicable — parameters per enclosure manufacturer (TBC) |

### 2.6.e · Explosion-Protection Calculations and Explanations (if applicable)

Flame-path (joint gap, length) and free-internal-volume calculations specific to Ex d (Section 2.5, team recommendation) have not yet been performed, pending final enclosure geometry and ExCB confirmation.

Status: Not yet available — pending ExCB confirmation of the Ex d protection concept and final enclosure design.

#### 2.6.e.1 Gas access path and internal arrangement

Ambient gas enters through the stainless-steel mesh at the front of the enclosure. During the sampling cycle the internal DC fan draws gas through this mesh into the sensing chamber, where the eight MQ sensors are installed. Each MQ sensor has its own stainless-steel protective mesh around the sensing element, so the sampled gas passes two successive mesh structures before reaching the sensing element.

2.6.e.2 Identification and assessment of potential ignition sources

| Ref. | Potential ignition source | Mechanism | Exposure / existing feature | Assessment required | Status |
| --- | --- | --- | --- | --- | --- |
| IS-01 | MQ sensor internal heater | Hot surface | Two meshes between atmosphere and element | Measured sensor-body surface ≤78.4 °C (fan stalled, 24.8 °C ambient); confirm heater construction and any flame-arresting function from manufacturer data | Open |
| IS-02 | DC sampling fan | Winding heating, commutation, stalled rotor | Front mesh upstream of fan | Fan CIXIKEJI CX5010B5H, 5 V 0.23 A; body measured 36.3 °C; confirm locked-rotor behaviour | Open |
| IS-03 | Main PCB | Arc/spark, component failure, overheating | Gas accessibility of PCB volume TBC | Confirm whether the PCB volume is separated from the sensing chamber (see e.3.2) | Open |
| IS-04 | DC/DC converters | Semiconductor / inductor heating, switching fault | Depends on IS-03 | PTC, TVS and MOSFET input stage present; converter temperatures not yet measured | Open |
| IS-05 | 24 VDC terminals and connectors | Make/break spark, loose-contact heating | Depends on IS-03 | Confirm terminal location, rating, torque, creepage/clearance | Open |
| IS-06 | LED / audible alarm module | Switching, heating, fault | Separate beacon housing | Obtain module construction/certification data | Open |
| IS-07 | ESP32-S3 / LoRa / digital electronics | Overheating, electrical fault | Depends on IS-03 | Assess maximum component temperatures | Open |
| IS-08 | RF transmitter / antenna | RF energy | External antenna | Max. 22 dBm into a 3 dBi antenna; assess against IEC 60079-0 RF threshold | Open |
| IS-09 | Power inductors | Surface heating | Depends on IS-03 | Measure worst-case ΔT | Open |
| IS-10 | MOSFETs, diodes, protection devices | Local overheating on overload/surge | Depends on IS-03 | Verify ratings and protective-device coordination (F1/F2 voltage rating) | Open |
| IS-11 | Electrostatic charging of non-metallic parts | ESD | Gasket, antenna insulation, gland parts | Identify exposed non-metallic parts and dimensions (MAT register) | Open |
| IS-12 | Metallic enclosure and joints | Spark from poor bonding/impact | External atmosphere | Verify bonding continuity and grounding | Open |
| IS-13 | Cable entry / gland / blanking devices | Loss of protection | External atmosphere | Prototype gland is a general-purpose polymer type; finalize Ex d gland/plug list (EX-07, EX-08) | Open |
| IS-14 | Front stainless-steel mesh | Flame transmission path | Direct interface | Confirm construction; certificate/test if claimed as flame-arresting | Open |
| IS-15 | MQ sensor protective mesh | Flame transmission path | Around each sensing element | No flame-arresting claim without manufacturer documentation | Open |
| IS-16 | Mechanical impact / friction | Mechanical spark | Fan is the only moving part | Confirm fan clearance and retention | Open |
| IS-17 | Battery circuit | High fault current, cell heating | Battery path populated on main board | Decide exclusion/depopulation from the certified configuration | Decision required |

2.6.e.3 Explosion-protection concept and protective construction

e.3.1 Overall concept

Potential ignition sources are isolated from the external atmosphere, enclosed in a protective construction, or controlled so they cannot ignite it. The applicant proposes flameproof enclosure Ex d (marking proposal II 2G Ex db IIC T4 Gb), subject to ExCB confirmation.

e.3.2 Gas-accessible and protected volumes

The front sensing chamber is intentionally gas-accessible. Whether the main PCB is separated from it by a partition (PCB plate) or shares the same volume is still to be confirmed on the section drawing; the assessment of IS-03…IS-10 depends on this.

e.3.3 Front stainless-steel mesh

First protective element in the gas inlet; mechanical protection and, only where supported by certification, flame arresting.

e.3.4 MQ sensor protective construction

Each sensor's integrated stainless-steel mesh surrounds the sensing and heating structure; no Ex claim without manufacturer documentation.

e.3.5 DC fan

Located in the gas path behind the mesh; surface temperature, winding temperature, stalled rotor and commutation to be assessed.

e.3.6 Main electronics

Main PCB, converters, processing and RF electronics and connectors are treated as potential ignition sources; protection depends on e.3.2.

e.3.7 Flame-transmission paths

Front gas inlet, cover joint, threaded interfaces, cable entries, antenna interface, alarm-module connection and any other wall penetration are to be identified and controlled; gap/length/volume values to be supplied by the enclosure manufacturer.

e.3.8 Cable entries, antenna, alarm module

All penetrations use components compatible with the final concept, gas group IIC, temperature range, IP66 and the M20×1.5 thread.

e.3.9 Grounding and electrostatic protection

The metallic enclosure is bonded to protective earth through the external grounding stud; conductive parts keep electrical continuity.

e.3.10 Temperature protection

Surface temperatures of sensors, fan, converters, enclosure and alarm module are evaluated at maximum load and fault conditions (2.6.f); T-class assigned only after testing.

e.3.11 Fault and abnormal conditions

Fan stall (measured), blocked airflow, heater abnormality, converter failure, short circuit, loose terminal and maximum ambient are to be covered.

### 2.6.f · Temperature Group Calculation (Hottest-Point Temperature Estimation)

A preliminary internal surface-temperature measurement has been carried out on a GLD V2 unit (eight thermocouple channels plus ambient; normal operation, maximum load, and fan stalled; 180 minutes each at approximately 24–25°C ambient). Maximum values, extrapolated to the +60°C maximum rated ambient:

| Condition | MQ sensor body | PCB | External enclosure |
| --- | --- | --- | --- |
| Normal operation | 44.2 → ≈79°C | 52.4 → ≈88°C | 32.1 → ≈67°C |
| Maximum load | 55.0 → ≈91°C | 58.3 → ≈94°C | 34.8 → ≈71°C |
| Fan stalled (malfunction) | 78.4 → ≈114°C | 63.4 → ≈99°C | 43.9 → ≈79°C |

The highest extrapolated value (≈114°C) is below the effective T4 limit of 130°C (135°C less the 5 K margin under IEC 60079-0). Under the recommended Ex d concept, the external enclosure surface (≈79°C) governs the temperature class. Open items: DC/DC converter, power inductor and MOSFET/diode not yet measured; alarm/transmit condition not measured; maximum-load PCB reading not yet at thermal equilibrium; test metadata (date, unit, thermocouple type, supply voltage) not yet recorded; confirmation at maximum ambient requires ExCB type testing. Full detail in the companion technical document.

Status: Partially available — preliminary internal measurement; indicates T4 is achievable with margin.

#### 2.6.f.1 Temperature-class evaluation structure (revision guide f.1–f.12)

f.1 Proposed temperature class

T4 (≤135 °C; effective limit 130 °C with the 5 K margin). Final assignment only after the hot-spot measurements, extrapolation to maximum ambient and fault conditions are complete.

f.2 Maximum ambient temperature

+60 °C, the rated maximum of the operating range −20 °C to +60 °C. Note: an assessment at +85 °C would raise the fan-stalled sensor-body value to about 139 °C, above the T4 limit; +85 °C is therefore not the rated ambient.

f.3 Candidate hottest points

| Ref. | Component / location | Reason | Measurement status |
| --- | --- | --- | --- |
| T-01 | MQ sensor heater / sensor head | Heater active during operation | Measured (TC-1) |
| T-02 | DC fan motor/electronics | Sampling and stalled condition | Measured (TC-2) |
| T-03 | 5 V converter (U36 LMR51450) | Supplies MQ heaters | Not measured (TC-3 planned) |
| T-04 | 3.3 V regulator (U43/U42) | MCU supply | Not measured |
| T-05 | Boost/buck converter ICs | Switching losses | Not measured (TC-3 planned) |
| T-06 | Power inductors | Copper/core losses | Not measured (TC-4 planned) |
| T-07 | Input MOSFET / protection devices | Conduction/fault heating | Not measured (TC-5 planned) |
| T-08 | ESP32-S3 | Processing / AI inference | Covered by PCB hottest area (TC-6) |
| T-09 | LoRa module | RF transmission | Not measured at maximum TX |
| T-10 | LED/buzzer alarm module | Alarm active | Not measured |
| T-11 | Main enclosure external surface | Governs T-class under Ex d | Measured (TC-7/TC-8) |
| T-12 | Front mesh / cover near sensor chamber | Heater and airflow | Not measured |

f.4 Worst-case operating condition

maximum permitted input supply voltage

maximum ambient temperature (+60 °C)

all eight MQ heaters energized

ESP32-S3 performing AI inference

LoRa transmitting at the maximum approved power (22 dBm)

local alarm active

DC fan operating according to the production sampling cycle

f.5 Fan failure / airflow loss

Measured: fan-stalled run of 180 min. Sensor body rose from 44.2 °C (normal) to 78.4 °C, PCB from 52.4 °C to 63.4 °C, external enclosure from 32.1 °C to 43.9 °C — the fan-stalled condition is the governing case so far.

f.6 Blocked or restricted mesh

Not tested. To be included where it affects internal cooling.

f.7 Electrical fault conditions

Regulator fault, heater overvoltage, shorted load, abnormal high current and component overload: not tested; to be assessed by test or by circuit-protection analysis (EL-01…EL-07).

f.8 Measurement method

Thermocouples at eight points plus ambient reference, logged for 180 minutes per condition. To be recorded for each point: ambient, input voltage, current, operating mode, fan/alarm/RF state, stabilization time and temperature. Thermocouple type and calibration: not yet recorded.

f.9 Thermal stabilization

Criterion ≤2 K/h. The PCB reading under maximum load was still rising in the final hour (46.4 → 55.8 → 58.3 °C) and must be repeated until stable.

f.10 Temperature correction / extrapolation

T(Ta,max) = Ta,max + (T,measured − T,amb,test). Example from the fan-stalled run: 60 + (78.4 − 24.8) ≈ 114 °C. Use of this method is subject to ExCB acceptance; confirmation at maximum ambient is part of type testing.

f.11 Results table

| Ref. | Point | Condition | Test ambient | Measured | Rise | Est. at +60 °C | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| T-01 | MQ sensor body | Normal | 24.9 °C | 44.2 °C | 19.3 K | ≈79 °C | Measured |
| T-01 | MQ sensor body | Maximum load | 24.2 °C | 55.0 °C | 30.8 K | ≈91 °C | Measured |
| T-01 | MQ sensor body | Fan stalled | 24.8 °C | 78.4 °C | 53.6 K | ≈114 °C | Measured |
| T-02 | Fan motor | Normal / stalled | 24.9 / 24.8 °C | 36.3 °C | 11.4 / 11.5 K | ≈71 °C | Measured |
| T-03 | 5 V converter IC | Maximum load | — | — | — | — | Not measured |
| T-04 | Power inductor | Maximum load | — | — | — | — | Not measured |
| T-05 | PCB hottest area (incl. ESP32-S3) | Fan stalled | 24.8 °C | 63.4 °C | 38.6 K | ≈99 °C | Measured |
| T-05 | PCB hottest area | Maximum load | 24.2 °C | 58.3 °C | 34.1 K | ≈94 °C | Preliminary — not stable |
| T-06 | LoRa module | Maximum TX | — | — | — | — | Not measured |
| T-07 | Alarm module | Alarm active | — | — | — | — | Not measured |
| T-08 | External enclosure | Fan stalled | 24.8 °C | 43.9 °C | 19.1 K | ≈79 °C | Preliminary — check TC-7 contact |

f.12 Preliminary temperature-class statement

The GLD is proposed for temperature class T4. The final classification remains open pending the remaining hot-spot measurements (T-03, T-04, T-06, T-07, T-09, T-10, T-12), thermal stabilization at maximum load, maximum-ambient confirmation and the relevant fault conditions. The highest extrapolated value so far, ≈114 °C, is below the effective T4 limit of 130 °C.

### 2.6.g · Usage and Installation Instructions (Draft)

Safety warnings:

Only trained and authorized personnel may install, commission, operate, inspect, or maintain the detector.

Isolate and verify the 24 VDC supply before opening the enclosure or changing wiring. Apply the site lockout/tagout procedure.

Use only the approved production power arrangement: 24 VDC.

Perform alarm tests with appropriate site controls and hearing protection where audible devices are connected.

Perform sensor nulling only in confirmed clean air. Accept the operation only when all eight sensors report a complete 8/8 result; otherwise correct the cause and repeat.

Do not drill, weld, enlarge cable entries, bypass protective devices, or modify the enclosure without written engineering and Ex approval.

Remove the unit from service if there is visible damage, water ingress, an unresolved sensor fault, failed alarm proof, or a loss of protective-earth continuity.

Detailed installation requirements:

Cable entry method: use only cable glands, adaptors, blanking elements, and seals approved for the final protection concept, certificate, thread form, IP rating, cable outer diameter, and installation temperature. Do not create new entries, enlarge an entry, leave unused entries open, or substitute a general-purpose gland. Maintain the manufacturer sealing/clamping sequence and record the installed component part number and certificate reference.

Torque requirements: torque values shall comply with the approved enclosure, terminal, cable-gland, and blanking-plug manufacturer instructions (specific torque figures not yet available, dependent on final hardware selection).

Grounding requirements: connect the enclosure to the approved protective-earth system using the designated grounding terminal; verify continuity before energization.

Cleaning requirements: isolate the equipment per site procedure and use a soft lint-free cloth lightly dampened with clean water or an approved mild cleaner; do not use abrasives, aggressive solvents, dry compressed air, high-pressure spray, or tools that can damage the enclosure, gasket, label, antenna, or cable entries; keep liquid out of all openings.

Operating instructions and maintenance requirements:

Before each operating period, check power, enclosure condition, diagnostic status, communication status, and the absence of active faults. Treat an alarm as a real process-safety event until the site response procedure establishes otherwise. Do not change firmware, model files, alarm thresholds, or calibration parameters from the field without controlled approval and a documented rollback path.

Status: Partially available (draft) — not yet consolidated into an Ex-specific installation and operation manual for ExCB review; also feeds the standalone Instruction_Manual_GLD deliverable.

#### 2.6.g.1 Usage and installation instructions — complete structure (g.2 … g.20)

The following sections complete the draft instructions above in the structure of the revision guide (g.1 Safety warnings is given in 2.6.g.a). Values not yet defined are stated as such in the notes of each section.

g.2 Permitted installation environment

The Gas Leak Detector (GLD) is intended for fixed installation in hazardous areas associated with oil and gas refinery facilities, including process units, tank farms, pipe racks, and storage/loading-unloading areas in which flammable gas atmospheres may be present.

The current certification target is Equipment Group II, Category 2G / EPL Gb, for installation in Zone 1 gas hazardous areas, subject to confirmation by the ExCB. The proposed gas group is IIC, and the proposed temperature class is T4. Final confirmation of the gas group and temperature class shall be based on the completed explosion-protection assessment, applicable component documentation, and the results of the required temperature evaluation and testing.

The GLD is intended to operate within an ambient temperature range of −20 °C to +60 °C and an operating relative-humidity range of 5–95% RH, non-condensing, as defined in the current product specification.

Installation shall be limited to locations and environmental conditions that are within the certified equipment marking and the conditions of use stated in the final certificate and approved instruction manual. The equipment shall not be installed in a hazardous area having a more severe gas group, temperature classification, ambient-temperature range, or other environmental condition than those covered by the final certification.

The GLD is intended for gas hazardous-area service only within the current certification scope. Installation in combustible-dust hazardous areas is excluded unless such use is separately assessed, tested, and approved as part of the certification scope.

The detector shall be installed in a position that permits free access of the surrounding atmosphere to the front gas-sampling inlet. The stainless-steel gas-entry mesh and the associated sampling path shall not be obstructed, covered, sealed, painted, or otherwise modified in a manner that could impair gas access to the sensing section.

The unit shall be mounted as a fixed installation using the approved mounting arrangement and in accordance with the orientation and mechanical details specified in the controlled drawings. Adequate clearance shall be maintained around the gas-sampling inlet, cable entries, antenna interface, and enclosure joints to permit correct operation, inspection, and maintenance.

The installation location shall protect the equipment from mechanical damage, unauthorized modification, excessive contamination, and environmental conditions beyond the approved operating limits. Cable entries, grounding connections, antenna interfaces, and other enclosure penetrations shall be installed only in accordance with the approved product configuration and the applicable installation instructions.

g.3 Mechanical mounting

The Gas Leak Detector (GLD) shall be installed as a fixed-mounted device using the approved mechanical mounting bracket provided for the product. The mounting arrangement shall correspond to the configuration defined in the controlled mechanical drawings and shall not be modified without engineering review and approval.

The GLD is intended to be attached to existing plant structures by means of the approved bracket arrangement. Installation shall not require drilling, welding, machining, or other permanent modification of the supporting plant structure unless such modification is separately approved by the site owner and is outside the certified GLD construction.

The mounting bracket, fasteners, clamps, and associated hardware shall be of the specified type, material, size, and mechanical strength defined in the approved production drawings and bill of materials. Substitution of mounting components that may affect the mechanical integrity or certified configuration of the equipment shall not be permitted without controlled engineering approval.

The detector shall be mounted securely so that it cannot become loose, rotate unintentionally, or move under normal service conditions, vibration, maintenance activity, or foreseeable external mechanical loading. All mounting fasteners shall be tightened in accordance with the torque values specified in the approved assembly or installation documentation.

The installed orientation shall conform to the orientation specified in the controlled mechanical drawings. The installation shall maintain unobstructed access to the gas-sampling inlet, front stainless-steel mesh, antenna interface, cable-entry points, grounding connection, and enclosure joints required for normal operation, inspection, and maintenance.

The gas-sampling inlet shall not be positioned directly against a wall, pipe, mounting plate, structural member, or other obstruction that could restrict the flow of ambient gas into the sensing area. Adequate clearance around the sampling inlet shall be maintained in accordance with the approved installation drawing.

The mounting location shall be selected so that the enclosure and its associated components are protected from foreseeable mechanical impact, excessive vibration, deformation, and other mechanical stresses that could impair the integrity of the enclosure, gas-entry assembly, cable entry, antenna interface, alarm module, or grounding connection.

After installation, the mounting assembly shall be visually inspected to confirm correct orientation, secure fastening, absence of enclosure deformation, and unobstructed access to the sensing inlet and service interfaces.

g.4 Gas inlet / sensing-area requirements

The Gas Leak Detector (GLD) shall be installed and maintained so that the gas-sampling inlet and sensing area remain continuously accessible to the surrounding atmosphere during normal operation.

Ambient gas enters the detector through the stainless-steel mesh located at the front of the enclosure. During the sampling cycle, the internal DC fan assists the movement of ambient gas through the front mesh and into the sensing area. Eight MQ-series gas sensors are installed within this sensing area. Each MQ sensor incorporates an additional stainless-steel protective mesh surrounding its sensing element.

The front gas-entry mesh, sampling opening, and associated airflow path shall remain unobstructed. The gas inlet shall not be covered, sealed, painted, taped, coated, or otherwise modified in a manner that could restrict or alter the intended gas flow to the sensing area. No additional filter, membrane, screen, protective cover, or other material shall be installed over the gas inlet unless specifically approved as part of the certified product configuration.

The detector shall be installed with sufficient clearance around the gas-sampling inlet to allow free circulation of the surrounding atmosphere. The front sensing inlet shall not be positioned directly against a wall, pipe, mounting plate, structural member, insulation, or other object that could restrict gas entry or interfere with the intended sampling airflow. The required minimum clearance, where applicable, shall be defined in the controlled installation drawing.

The stainless-steel front mesh shall be maintained in its approved production configuration. It shall not be removed, replaced, drilled, enlarged, reshaped, or substituted with a different mesh or protective element unless the change has been evaluated and approved through the applicable engineering and Ex certification change-control process.

The integrated protective mesh of each MQ sensor shall also remain intact and free from mechanical damage, deformation, contamination, or blockage. MQ sensors with damaged or missing protective mesh shall not be used.

The sensing area shall be kept free from excessive dust, deposits, oil, paint, corrosion products, insects, or other contamination that could obstruct the front mesh, fan, individual sensor meshes, or the intended gas-sampling path. Cleaning shall be performed only in accordance with the approved maintenance instructions and shall not damage the mesh, fan, sensor surfaces, enclosure sealing features, or other Ex-critical components.

The DC sampling fan shall be installed in the approved position and orientation behind the front stainless-steel mesh. The airflow direction, fan mounting, electrical connection, and mechanical clearance shall conform to the controlled assembly drawing. The fan shall not be substituted with a different model without controlled engineering review and approval.

The gas-sampling path shall be considered part of the controlled product construction. Any modification that changes the geometry, mesh arrangement, fan position, airflow path, sensing-chamber volume, or position of the MQ sensors shall require evaluation before implementation.

During inspection and maintenance, the front mesh, fan, sensing area, and individual MQ sensor meshes shall be checked for blockage, damage, contamination, corrosion, or other conditions that could affect gas access or the approved protection concept. If any such condition is detected, the detector shall be removed from normal service until the condition has been corrected and the unit has been verified as suitable for operation.

Note: The minimum clearance in front of the gas inlet, the approved fan orientation, and any dimensional requirements of the sensing chamber shall be specified in the controlled mechanical and installation drawings.

g.5 Electrical supply

The Gas Leak Detector (GLD) shall be supplied from a 24 VDC power source in accordance with the approved product configuration. The current production configuration is intended for continuous operation from a 24 VDC supply. The product documentation presently identifies the 24 VDC supply as being derived from an AC/DC adapter connected to the site electrical supply.

The GLD internal power system converts the 24 VDC input to the required internal operating voltages. A 5 VDC rail is used for the MQ sensor/heater circuitry, while a 3.3 VDC rail is used for the ESP32-S3 logic and other applicable low-voltage electronics.

The currently documented maximum power consumption is 8 W at 24 VDC. The corresponding calculated input current at this operating point is approximately 0.33 A. This value is derived from the measured maximum power consumption divided by the nominal 24 VDC input voltage.

The 24 VDC input incorporates electrical protection features including resettable PTC devices, a 33 V TVS device on the 24 V input, and P-channel MOSFET input stages. These circuits are intended to provide protection against overcurrent, transient overvoltage, and reverse-polarity or other abnormal input conditions as defined by the final approved electrical design.

The polarity of the 24 VDC supply shall be verified before energizing the equipment. Power shall be connected only through the designated power-input terminals or connector defined in the approved wiring and terminal drawings. The power-input connection shall not be modified, bypassed, or connected to an alternative supply arrangement that is not included in the approved product configuration.

The equipment shall be de-energized before opening the enclosure, connecting or disconnecting field wiring, or carrying out maintenance activities that require access to the internal electrical system. The applicable site isolation and lockout/tagout procedures shall be followed before such work is performed.

The 24 VDC supply, associated field wiring, cable entry, terminals, protective-earth connection, and upstream protective devices shall be installed in accordance with the approved product documentation and applicable site electrical requirements.

Any alternative power source, power-conversion arrangement, battery supply, or other electrical configuration that is not included in the approved production configuration shall not be used unless it has been separately evaluated and approved through the applicable engineering and Ex certification change-control process.

Note: The permitted 24 VDC input-voltage tolerance, maximum startup/inrush current, required upstream fuse or circuit-breaker rating, and detailed terminal/wire-size requirements shall be defined in the final approved electrical specification and installation drawing.

g.6 Cable entry

The Gas Leak Detector (GLD) shall use only cable-entry devices that are approved for the final certified product configuration. Cable glands, adaptors, blanking elements, seals, and associated entry components shall be compatible with the final explosion-protection concept, certificate requirements, thread form, ingress-protection rating, cable outer diameter, and permitted installation temperature range.

The cable-entry arrangement shall be installed exactly as defined in the approved mechanical drawings and installation instructions. No cable-entry opening shall be drilled, enlarged, re-machined, re-threaded, or otherwise modified in the field. Unused entries shall not be left open and shall be closed using the approved blanking element specified for the certified configuration.

The cable gland or entry device shall be selected for the actual cable type and cable outer diameter used in the installation. The manufacturer’s specified sealing, clamping, and assembly sequence shall be maintained. The installed cable shall be properly retained so that mechanical loading is not transferred to the internal electrical terminals.

All threaded cable-entry components shall be fully engaged in accordance with the minimum thread-engagement requirements defined by the approved enclosure and cable-entry configuration. Where a locking element is specified, it shall be applied only after the required thread engagement has been achieved.

Tightening of cable glands, adaptors, blanking elements, and related fittings shall comply with the torque values specified by the approved component manufacturer and the controlled installation documentation. Excessive tightening that could damage threads, seals, or enclosure parts shall be avoided, and insufficient tightening that could compromise sealing or mechanical retention shall not be permitted.

During installation and maintenance, the cable-entry components shall be inspected for correct part number, thread compatibility, sealing-element condition, proper cable clamping, complete thread engagement, and absence of mechanical damage, corrosion, cracking, or deformation. Components showing damage or deterioration shall not be reused unless specifically permitted by the approved component instructions.

The installed cable-entry component shall be traceable to the approved product configuration. The manufacturer, model or part number, thread type, applicable certificate reference, and installation details shall be recorded where required by the production or installation documentation. The current draft already requires recording the installed component part number and certificate reference.

Any substitution of a cable gland, adaptor, blanking plug, sealing element, or other cable-entry component shall be subject to controlled engineering review before implementation. A general-purpose cable gland or an entry component not covered by the approved product configuration shall not be used.

Note: The final cable-entry device manufacturer, model, thread designation, cable-diameter range, temperature range, ingress-protection rating, applicable Ex marking, certificate reference, and installation torque shall be confirmed in the controlled BOM and installation drawing before certification documentation is finalized.

g.7 Grounding and bonding

The metallic enclosure of the Gas Leak Detector (GLD) shall be connected to the site protective-earth system using the designated external grounding stud provided on the enclosure. The grounding point shall be clearly identified by the protective-earth symbol and shall remain accessible for installation, inspection, and maintenance.

The protective-earth connection shall be installed using the approved grounding hardware, including the specified grounding stud, terminal lug, washers, fasteners, and other associated components defined in the controlled mechanical drawing and bill of materials.

All conductive enclosure parts that are required to be electrically bonded shall maintain reliable electrical continuity so that hazardous potential differences and electrostatic charge accumulation are minimized. The bonding arrangement between enclosure parts shall not be modified, insulated, painted over at the contact interface, or otherwise altered in a manner that could impair electrical continuity.

Before the GLD is energized, the protective-earth connection and bonding continuity shall be verified. The continuity check shall confirm that the external grounding stud is electrically connected to the applicable conductive enclosure parts in accordance with the approved design.

The grounding conductor, terminal lug, and grounding hardware shall be mechanically secure and protected against loosening, corrosion, and accidental disconnection. Tightening shall be carried out in accordance with the torque values specified in the approved assembly and installation documentation.

Where surface coatings, anodizing, paint, oxidation, or other insulating layers are present, the grounding and bonding contact surfaces shall be prepared or constructed so that the required electrical continuity is maintained. No unapproved coating, sealant, adhesive, or insulating material shall be introduced between designated bonding surfaces.

The protective-earth connection shall not be used as a mechanical support point for unrelated cables or equipment. The grounding arrangement shall be reserved for the intended protective-earth and bonding function defined by the approved product design.

During periodic inspection and maintenance, the grounding stud, lug, fasteners, bonding points, and accessible conductive connections shall be checked for looseness, corrosion, contamination, mechanical damage, or any other condition that could impair electrical continuity. Any defective grounding or bonding connection shall be corrected before the equipment is returned to service.

Any modification to the grounding stud, bonding arrangement, enclosure contact surfaces, grounding conductor arrangement, or associated hardware shall be subject to controlled engineering review before implementation.

Note: The grounding-conductor size, terminal-lug type, fastening torque, bonding arrangement, and continuity acceptance criteria shall be defined in the final controlled grounding drawing and installation documentation.

posisi grounding stud,

diameter/ukuran stud,

urutan washer–lug–nut,

material hardware,

simbol PE,

nilai torque,

acceptance limit continuity.

g.8 Antenna installation

The Gas Leak Detector (GLD) uses an external omnidirectional antenna connected to the internal LoRa communication module through the approved RF interface. The production configuration specified in this documentation uses a 3 dBi external antenna with an SMA interface

The antenna shall be installed only in the position and orientation defined in the approved mechanical and assembly drawings. For the current production configuration, the antenna interface is located at the upper part of the base enclosure. Earlier side-mounted antenna arrangements shown in superseded photographs or drawings shall not be used for the certified production configuration.

The antenna, SMA connector, internal RF cable, U.FL connection where applicable, mounting hardware, sealing elements, and enclosure penetration shall correspond to the approved production configuration. Substitution of the antenna, connector, RF cable, bulkhead fitting, or associated mounting parts shall not be permitted without controlled engineering review and approval.

The antenna shall be mounted securely so that vibration, accidental movement, or handling does not loosen the RF connector or compromise the enclosure interface. The SMA connection shall be fully engaged and tightened in accordance with the approved component and assembly instructions. Excessive torque that could damage the connector, enclosure, sealing element, or internal RF cable shall be avoided.

The antenna mounting arrangement shall maintain the required enclosure sealing and mechanical integrity. Any gasket, O-ring, sealing washer, bulkhead fitting, or other sealing component associated with the antenna penetration shall be installed in the correct sequence and orientation defined in the controlled drawing.

The antenna installation shall not introduce any unapproved opening, gap, modification, drilling, or additional penetration through the enclosure wall. The antenna interface shall not be enlarged, re-machined, re-threaded, or otherwise modified in the field.

The external antenna shall be positioned so that it is not mechanically obstructed or excessively loaded by nearby plant structures, piping, cable trays, or other equipment. The installation arrangement shall minimize the risk of mechanical impact, bending, loosening, or damage to the antenna and its enclosure interface.

The antenna shall not be used as a lifting point, support point, handle, or attachment point for other equipment or cables.

During installation and periodic inspection, the antenna and RF interface shall be checked for secure fastening, corrosion, cracking, deformation, damaged insulation, loose connectors, damaged sealing components, and other conditions that could affect RF performance, enclosure integrity, or the approved Ex configuration. Any damaged component shall be replaced only with an approved equivalent specified in the controlled BOM.

The antenna type, gain, RF connector, cable assembly, and mounting arrangement shall remain consistent with the approved radio configuration. Any change in antenna type, antenna gain, cable length, RF connector arrangement, or mounting position shall be subject to engineering review to determine whether it affects the approved RF and explosion-protection configuration.

Note: The approved antenna manufacturer, model, maximum antenna gain, SMA connector part number, RF cable type and length, mounting torque, sealing arrangement, and final installation position shall be specified in the controlled BOM and mechanical drawing.

g.9 Commissioning

Commissioning of the Gas Leak Detector (GLD) shall be performed only after the mechanical installation, electrical connection, cable-entry installation, grounding/bonding connection, antenna installation, and enclosure assembly have been completed in accordance with the approved installation documentation.

Commissioning shall be carried out only by trained and authorized personnel. Before energizing the equipment, the installer shall verify that the GLD is installed in the permitted hazardous-area environment and that the installation corresponds to the approved product configuration.

Prior to initial energization, the following items shall be verified:

the enclosure is correctly assembled and free from visible mechanical damage, corrosion, deformation, or contamination;

the front stainless-steel gas-entry mesh is clean, intact, correctly installed, and unobstructed;

the gas-sampling inlet has adequate clearance from surrounding structures;

the DC sampling fan is correctly installed and mechanically unobstructed;

all cable-entry devices, adaptors, blanking elements, and seals are correctly installed and securely tightened;

the 24 VDC input wiring is connected with the correct polarity;

the protective-earth and bonding connections are securely installed and electrical continuity has been verified;

the antenna and RF interface are correctly installed and securely fastened; and

all enclosure covers, threaded joints, fasteners, and locking features are correctly engaged in accordance with the approved assembly instructions.

After completion of the pre-energization inspection, the GLD may be energized from the approved 24 VDC supply. The supply voltage, polarity, and equipment current consumption shall be checked for conformity with the approved electrical specification. The current production configuration identified in the draft is based on continuous 24 VDC operation.

Following power-up, the GLD shall be allowed to complete its normal start-up and sensor stabilization sequence in accordance with the approved operating procedure. The operator shall verify that the processing unit, sensor-acquisition system, sampling fan, alarm module, and communication functions initialize without an active fault condition.

A sensor baseline or nulling procedure shall be performed only in confirmed clean air. The nulling procedure shall be considered acceptable only when all eight MQ sensor channels report a complete 8/8 successful result. If the nulling sequence is incomplete or any sensor channel reports a fault, the cause shall be investigated and corrected before the detector is placed into normal service.

The sampling function shall be verified by confirming that the internal DC fan operates during the applicable sampling cycle and that no physical obstruction exists in the gas-inlet or sensing path.

The local alarm function shall be tested to confirm correct operation of the integrated visual and audible alarm. Where a controlled gas test is used, the test shall be performed in accordance with the approved site procedure and with appropriate safety precautions.

The LoRa communication function shall be verified by confirming successful transmission of status and/or alarm information from the GLD to the designated cluster head or gateway. Any required wired diagnostic or Modbus communication interface may also be verified as part of the commissioning procedure where applicable.

Before the GLD is released for normal operation, the operator shall confirm that:

no unresolved sensor or system fault is present;

all eight sensor channels are operational;

the local alarm function has been verified;

the LoRa communication link is operational;

the grounding connection is intact;

the gas-sampling inlet remains unobstructed; and

the enclosure, cable entries, antenna interface, and other Ex-critical mechanical features remain correctly assembled.

If any required commissioning check is unsuccessful, the GLD shall not be placed into normal detection service until the cause has been identified, corrected, and the affected commissioning checks have been repeated successfully.

The commissioning results shall be recorded and shall include, as applicable, the product model and serial number, hardware revision, firmware version, AI model/configuration version, date of commissioning, supply voltage, nulling result, alarm-test result, communication-test result, grounding-continuity result, and identification of the responsible commissioning personnel.

Note: The required sensor stabilization/warm-up time, permitted 24 VDC input range, acceptance limits for grounding continuity, maximum commissioning current, and detailed gas-test procedure shall be defined in the final approved operating and commissioning procedure.

g.10 Operating modes

The Gas Leak Detector (GLD) provides defined operating modes for normal gas-detection service, engineering/data-collection activities, and sensor baseline adjustment. Each operating mode shall be used only for its intended purpose and in accordance with the approved operating procedure.

Inference Mode

Inference Mode is the normal production operating mode of the GLD. In this mode, the detector acquires data from the eight MQ-series gas sensor channels, processes the sensor data using the embedded processing and gas-classification functions, evaluates the configured alarm criteria, and performs the normal alarm and communication functions.

Inference Mode shall be used after installation and commissioning have been successfully completed. The detector shall remain in this mode during normal gas-monitoring service unless an authorized maintenance, calibration, or engineering activity requires another approved mode.

Dataset Mode

Dataset Mode is intended for engineering and controlled data-collection activities. This mode may be used to acquire sensor data for development, verification, characterization, or other approved engineering purposes.

Dataset Mode is not a substitute for normal approved gas-detection operation and shall not be relied upon as the production operating mode while the detector is being used for normal hazardous-area gas monitoring. This limitation is already stated in the current draft.

Use of Dataset Mode shall be restricted to authorized personnel and shall be performed under a controlled procedure. Any engineering activity that changes firmware, AI model files, alarm thresholds, calibration parameters, or other controlled configuration items shall be subject to the applicable engineering change-control process.

Nulling Mode

Nulling Mode is intended for sensor baseline adjustment before or during approved commissioning and maintenance activities. The nulling procedure shall be performed only in confirmed clean air so that the baseline values of the eight MQ sensor channels are not influenced by the presence of target or interfering gases.

The nulling operation shall be considered acceptable only when all eight sensor channels complete the procedure successfully and report a complete 8/8 result. If the nulling process is incomplete or one or more channels do not pass, the cause shall be investigated and corrected before the detector is returned to normal service. The current draft already specifies clean-air operation and the complete 8/8 requirement.

After completion of Nulling Mode, the detector shall be returned to Inference Mode before being released for normal gas-detection operation.

Operating-Mode Control

Selection or change of operating mode shall be restricted to authorized operation, commissioning, maintenance, or engineering procedures. Operating-mode changes shall not bypass or disable any required safety, alarm, diagnostic, or configuration-control function of the approved product.

The active operating mode shall be identifiable through the approved diagnostic or service interface where applicable. Before the detector is released for normal service, the responsible personnel shall verify that the GLD has been returned to Inference Mode and that no unresolved fault or incomplete maintenance condition remains.

g.11 Normal operation

Following successful installation, commissioning, and completion of the required sensor nulling procedure, the Gas Leak Detector (GLD) shall be placed in Inference Mode for normal gas-monitoring operation. Inference Mode is the approved production operating mode of the detector.

Before each operating period, the responsible personnel shall verify that the GLD is energized correctly, that the enclosure and external components show no visible damage, that the diagnostic status is normal, that the required communication link is available, and that no unresolved fault condition is present.

During normal operation, the GLD monitors the surrounding atmosphere using the eight MQ-series gas sensor channels. Sensor data are acquired and processed locally by the ESP32-S3 processing unit, and the embedded gas-classification function evaluates the multi-sensor response pattern. The principal detection and classification functions are performed locally within the GLD and therefore do not depend on a continuous connection to a central server.

During the applicable sampling cycle, ambient gas is admitted into the sensing area through the front stainless-steel mesh. The internal sampling arrangement assists the exposure of the eight MQ-series sensors to the surrounding atmosphere. Each MQ sensor provides an electrical response that is acquired by the detector electronics for subsequent processing and classification.

When the processed sensor response satisfies the configured alarm criteria, the GLD shall activate the integrated local visual and audible alarm and transmit the alarm information through the LoRa communication network to the associated monitoring system. The current product description defines these local and remote alarm paths as parallel alarm functions.

Any gas alarm generated during normal operation shall be treated as a genuine process-safety event until the applicable site response procedure has established otherwise. The operator shall follow the site-specific alarm-response procedure and shall not disregard, suppress, or bypass an active alarm without authorized investigation.

If the GLD reports a sensor, system, or diagnostic fault, or if the required communication link is lost, the responsible operator shall be informed and the affected detection point shall be handled in accordance with the site impairment procedure. The detector shall not be considered fully available for normal monitoring while an unresolved fault condition remains.

During normal operation, the front gas-sampling mesh and sensing inlet shall remain unobstructed. The enclosure, antenna, cable entries, grounding connection, and other Ex-critical external features shall remain in their approved installed condition. No enclosure cover, cable-entry device, sensor assembly, mesh, antenna interface, or other controlled construction feature shall be removed or modified while the detector is in normal service.

Firmware, AI model files, alarm thresholds, calibration parameters, and other controlled configuration settings shall not be changed in the field unless the change has been authorized through the approved engineering and configuration-control process and a documented rollback method is available.

If visible enclosure damage, water ingress, an unresolved sensor fault, a failed alarm function, or loss of protective-earth continuity is detected, the GLD shall be removed from normal service until the condition has been corrected and the detector has been verified as suitable for continued operation.

Normal operation shall be resumed only after any maintenance or corrective action has been completed, the enclosure has been correctly reassembled, the required functional checks have been successfully performed, and the GLD has been returned to Inference Mode with no unresolved fault condition.

g.12 Alarm and fault handling

The Gas Leak Detector (GLD) shall provide alarm and fault indications in accordance with the approved production configuration and operating software. During normal operation, the detector continuously evaluates the processed response of the eight MQ-series gas sensor channels and the associated system diagnostic information.

When the configured alarm criteria indicate a gas-leak condition, the GLD shall initiate two parallel alarm paths:

1.  activation of the integrated local visual and audible alarm; and

2.  transmission of the alarm information through the LoRa communication network to the associated cluster head, gateway, or operator monitoring system, as applicable to the approved system configuration.

The local gas-detection and alarm-decision functions are performed within the GLD and shall not depend on a continuous connection to the remote monitoring system. Loss of the LoRa communication link shall therefore be treated as a communication or system-availability fault and shall not by itself be interpreted as confirmation that no hazardous gas condition exists.

Any gas alarm generated by the GLD shall be treated as a genuine process-safety event until the applicable site alarm-response procedure has established otherwise. The operator shall follow the site-specific emergency or gas-response procedure and shall not ignore, suppress, bypass, or disable an active gas alarm without authorized investigation.

Where the detector identifies a sensor fault, diagnostic fault, communication loss, power-related fault, or other condition that may impair the availability of the detection function, the responsible operator shall be notified and the affected detection point shall be placed into the applicable site impairment process until the fault has been corrected and the detector has been verified as suitable for continued operation.

A sensor-related fault shall be investigated whenever one or more of the eight MQ sensor channels fail the required diagnostic, baseline, or nulling criteria. A detector that cannot complete the required eight-channel sensor verification shall not be returned to normal gas-detection service until the cause has been identified and corrected. The current draft requires a complete 8/8 result for the nulling operation.

Fault investigation shall include, as applicable, verification of the 24 VDC power supply, internal sensor acquisition, sampling fan operation, local alarm module, LoRa communication, grounding continuity, enclosure condition, gas-entry path, and the status of the eight sensor channels.

If visible enclosure damage, water ingress, an unresolved sensor fault, a failed alarm proof test, or loss of protective-earth continuity is identified, the GLD shall be removed from service until the condition has been corrected and the detector has successfully completed the required inspection and functional verification.

Alarm or fault conditions shall not be cleared merely by power cycling, resetting, reconfiguring, or modifying the detector unless such action forms part of the approved operating or maintenance procedure. Any corrective action shall preserve the approved hardware, firmware, AI-model, alarm-threshold, calibration, and explosion-protection configuration.

Firmware, model files, alarm thresholds, and calibration parameters shall not be changed in the field as a means of suppressing an alarm or clearing a fault unless the change has been formally authorized through the controlled engineering and configuration-management process.

After a fault has been corrected, the GLD shall undergo the applicable functional checks before being returned to normal service. These checks shall include, as applicable, sensor-status verification, nulling in confirmed clean air, sampling-fan verification, local alarm verification, LoRa communication verification, grounding-continuity verification, and confirmation that no unresolved diagnostic condition remains.

The detector shall be returned to normal Inference Mode only after the required checks have been completed successfully and the responsible personnel have confirmed that the unit is suitable for continued operation.

| Condition | Required action |
| --- | --- |

| Gas alarm | Treat as a real process-safety event and follow the site gas-response procedure |
| --- | --- |
| Sensor fault | Investigate affected sensor channel(s); do not return to service until resolved |
| Nulling incomplete / not 8/8 | Correct the cause and repeat nulling in confirmed clean air |
| LoRa communication loss | Notify operator and place affected detection point in the site impairment process |
| Local alarm failure | Remove detector from normal service until repaired and verified |
| Grounding fault | De-energize/remove from service until continuity is restored |
| Water ingress / enclosure damage | Remove from service and inspect before recommissioning |
| Fault corrected | Repeat applicable functional checks and return to Inference Mode only after successful verification |

g.13 Cleaning

Cleaning of the Gas Leak Detector (GLD) shall be carried out only in accordance with the approved maintenance procedure and with the equipment isolated as required by the applicable site safety procedure.

For routine external cleaning, a soft, lint-free cloth lightly dampened with clean water or an approved mild cleaning agent shall be used. Cleaning shall be performed carefully so that the enclosure surface, labels, antenna, cable-entry components, seals, and other external parts are not damaged.

Abrasive materials, aggressive solvents, dry compressed air, high-pressure water spray, or tools that could damage the enclosure, gasket, label, antenna, cable entries, or other product components shall not be used. Liquid shall not be permitted to enter the equipment through openings, cable entries, enclosure joints, or other interfaces.

Particular attention shall be given to the condition of the front protective mesh and sensing area. The mesh shall remain free from visible deposits, contamination, corrosion products, or other material that could obstruct the gas-sampling path. The current maintenance schedule already requires periodic visual inspection of the sensor and protective mesh.

Cleaning of the sensing area shall be performed without deforming, puncturing, removing, or otherwise damaging the protective mesh or sensor assemblies. The integrated protective mesh of the MQ sensors shall not be mechanically altered during cleaning or maintenance.

The enclosure sealing surfaces, gasket, cable-entry seals, antenna interface, and other sealing components shall be visually inspected during cleaning for signs of deterioration, cracking, deformation, contamination, or mechanical damage. Any condition that could impair the enclosure or sealing performance shall be corrected before the detector is returned to service.

After cleaning, the GLD shall be inspected to confirm that no liquid, residue, foreign material, or cleaning agent remains in a manner that could interfere with the gas-sampling inlet, enclosure sealing, cable entry, antenna interface, grounding connection, or normal operation of the detector.

If cleaning reveals visible damage, water ingress, damaged sealing components, or other conditions that could affect safe operation, the GLD shall be removed from service until the condition has been corrected and the required inspection has been completed. The draft already requires removal from service in cases such as visible damage or water ingress.

Note: Any detailed cleaning method for the front stainless-steel mesh and the integrated MQ sensor mesh shall be defined only after the approved mesh construction and sensor manufacturer instructions have been confirmed.

g.14 Periodic inspection and maintenance

The Gas Leak Detector (GLD) shall be subject to periodic inspection and preventive maintenance throughout its service life in order to maintain the gas-detection function, enclosure integrity, electrical safety, and the explosion-protection characteristics of the approved product configuration.

Inspection and maintenance shall be performed only by trained and authorized personnel in accordance with the approved maintenance procedure and the applicable hazardous-area safety requirements of the installation site.

Before any maintenance activity requiring the enclosure to be opened, the GLD shall be isolated from the 24 VDC power supply and the applicable site lockout/tagout procedure shall be followed. The enclosure shall not be opened while an explosive atmosphere may be present unless the equipment has been de-energized and the area has been confirmed safe in accordance with the applicable site procedure.

Routine Operational Inspection — Daily / Each Shift

During each operating shift, or at the frequency defined by the site operating procedure, the operator shall verify the general operational condition of the detector. The inspection shall include:

verification that the GLD is correctly powered;

verification that no active sensor or system fault is present;

inspection of the external enclosure for visible damage, corrosion, deformation, or evidence of water ingress;

verification of the diagnostic status;

verification of the LoRa communication status; and

confirmation that the detector remains available for normal gas-monitoring operation.

These checks correspond to the daily/each-shift inspection currently defined in the maintenance schedule.

Three-Month Inspection

At intervals not exceeding three months, the gas-sampling inlet, front stainless-steel protective mesh, sensing area, and accessible sensor components shall be visually inspected. The inspection shall confirm that:

the front mesh is clean and unobstructed;

no excessive dust, oil, deposits, insects, corrosion products, or other contamination restricts the gas-sampling path;

the front mesh shows no deformation, puncture, mechanical damage, or corrosion;

the MQ sensor assemblies and their integrated protective meshes show no visible damage or contamination; and

the sensing inlet remains correctly positioned and unobstructed by surrounding structures.

Cleaning, where required, shall be carried out in accordance with Section g.13 and shall not damage or alter the approved mesh or sensing arrangement. The current draft specifies visual inspection of the sensors and protective mesh at three-month intervals.

Six-Month Inspection and Functional Verification

At intervals not exceeding six months, a more detailed functional and mechanical inspection shall be carried out. The following activities shall be performed as applicable:

verification of the local visual and audible alarm function;

verification of sensor status;

sensor nulling/baseline verification in confirmed clean air in accordance with the approved procedure;

inspection of the enclosure joints, threaded interfaces, and other explosion-protection-related surfaces for corrosion, contamination, wear, or mechanical damage;

inspection of the front mesh and gas-sampling path;

inspection of the DC sampling fan for correct operation and absence of obstruction or abnormal mechanical condition; and

verification that no unresolved diagnostic or communication fault remains.

The existing draft specifies alarm-function testing, sensor recalibration/nulling, and inspection of flame paths and cover threads at six-month intervals.

Where nulling is performed, it shall be carried out only in confirmed clean air. The procedure shall be accepted only when all eight sensor channels successfully complete the required 8/8 result. If the required result is not obtained, the cause shall be investigated and corrected before the GLD is returned to normal service.

Annual Comprehensive Inspection

At least once every twelve months, a comprehensive inspection of the complete GLD installation shall be performed. The inspection shall include, as applicable:

protective-earth and bonding continuity verification;

inspection of the external grounding stud, terminal lug, washers, and associated fastening hardware;

inspection and verification of cable glands, cable-entry devices, adaptors, blanking elements, and sealing components;

verification of terminal and applicable mechanical tightening conditions;

inspection of enclosure gaskets and seals for ageing, cracking, hardening, deformation, contamination, or other deterioration;

inspection of enclosure joints, threads, and safety-critical mechanical interfaces;

inspection of the antenna, SMA interface, sealing arrangement, and RF cable connection;

inspection of the alarm module and associated enclosure interface;

inspection of the front gas-entry mesh, sampling fan, sensing area, and MQ sensor assemblies;

verification of the local alarm function;

verification of LoRa communication;

review of sensor condition and remaining service life; and

confirmation that the hardware, firmware, AI model, and configuration correspond to the approved product configuration.

The current draft specifies annual grounding-continuity verification, cable-gland and terminal checks, gasket/seal assessment, and evaluation of sensor-element service life.

Maintenance of Ex-Critical Components

Components that form part of the approved explosion-protection construction shall not be repaired, replaced, modified, or substituted with unapproved alternatives. This includes, as applicable, the enclosure, cover, front stainless-steel mesh, cable-entry components, seals and gaskets, grounding hardware, antenna interface, alarm-module interface, sampling fan, and other components identified in the Ex-critical BOM.

If an Ex-critical component is damaged or requires replacement, only the approved replacement component specified in the controlled BOM or other formally approved documentation shall be used. Any proposed substitution shall be subject to controlled engineering review before implementation.

Conditions Requiring Removal from Service

The GLD shall be removed from normal service if inspection or maintenance identifies a condition that could impair the detector function or the approved protection concept, including visible enclosure damage, water ingress, an unresolved sensor fault, failed alarm verification, or loss of protective-earth continuity. These conditions are already identified in the current safety instructions.

The equipment shall not be returned to normal operation until the identified condition has been corrected and the applicable inspection, functional verification, and commissioning checks have been completed successfully.

Maintenance Records

Periodic inspection and maintenance activities shall be documented. Records shall include, as applicable, the equipment model and serial number, inspection date, operating condition, findings, maintenance activities performed, components replaced, nulling or functional-test results, grounding-continuity result, firmware/configuration identification where relevant, and identification of the personnel performing the work.

| Interval | Minimum activities |
| --- | --- |

| Daily / Each Shift | Power, enclosure condition, diagnostics, active faults, LoRa communication |
| --- | --- |
| Every 3 Months | Front mesh, MQ sensor area, contamination, blockage, visible sensor/mesh damage |
| Every 6 Months | Alarm test, nulling 8/8, fan operation, enclosure/thread/interface inspection |
| Annually | Comprehensive inspection, PE/bonding continuity, glands/terminals, gasket/seals, antenna interface, alarm, LoRa, sensor service-life review |

g.15 Sensor Nulling / Calibration

Sensor nulling shall be performed as part of commissioning and periodic maintenance in order to establish or verify the baseline condition of the eight MQ-series gas sensor channels.

The nulling procedure shall be carried out only in confirmed clean air, under conditions where the sensor responses are not influenced by the presence of target gases or other significant contaminants.

The nulling procedure shall be considered successful only when all eight sensor channels complete the procedure and report a complete 8/8 successful result. If one or more sensor channels fail to complete the nulling sequence, the detector shall not be returned to normal gas-detection operation until the cause has been identified and corrected.

Before nulling is initiated, the operator shall confirm that the GLD has completed its required start-up and sensor-stabilization sequence and that no unresolved system or sensor fault is present.

Nulling shall not be used to suppress, compensate for, or conceal a sensor fault, excessive sensor drift, contamination, abnormal sensor response, or other malfunction.

After successful completion of nulling, the GLD shall be returned to Inference Mode and the normal diagnostic status shall be verified before the detector is released for service.

Any calibration or gas-response verification beyond the nulling/baseline procedure shall be performed only in accordance with an approved calibration procedure using the applicable reference gas and test equipment. Calibration parameters shall not be changed in the field except through the controlled configuration-management process.

g.16 Inspection of Ex-Critical Parts

Components identified as Ex-critical shall be inspected at the specified maintenance intervals and whenever the equipment is opened, serviced, repaired, or otherwise disturbed.

Ex-critical parts include, as applicable, the enclosure body and cover, front stainless-steel mesh assembly, enclosure joints and threaded interfaces, gaskets and seals, cable-entry devices, adaptors, blanking elements, grounding and bonding hardware, antenna interface, alarm-module interface, sampling fan, MQ sensor assemblies, and other components identified in the controlled Ex-critical BOM.

Inspection shall confirm that each Ex-critical component:

corresponds to the approved manufacturer, model, material, and part number;

is correctly installed in the approved orientation and position;

shows no unacceptable corrosion, wear, cracking, deformation, contamination, or mechanical damage;

has not been drilled, machined, re-threaded, reshaped, coated, or otherwise modified without approval;

retains the specified sealing, fastening, thread engagement, or mechanical retention characteristics; and

remains consistent with the approved product drawings and bill of materials.

Enclosure threads, joints, mesh assemblies, gaskets, cable-entry components, and other interfaces affecting the protection concept shall be examined carefully for damage or deterioration before the equipment is returned to service.

Any Ex-critical part that is damaged, missing, incorrectly installed, or inconsistent with the approved configuration shall be treated as a non-conforming condition. The GLD shall not be returned to normal operation until the condition has been corrected and the affected protection-related features have been re-inspected.

g.17 Firmware / Configuration Control

The GLD firmware, embedded AI model, alarm parameters, communication settings, calibration data, and other controlled configuration items shall be maintained under formal configuration control.

Only firmware and AI-model versions that have been approved for the applicable GLD hardware revision shall be installed in production or field equipment.

The approved software configuration shall be identifiable by controlled version information, build identification, checksum/hash, or other traceable configuration identifier as defined by the project configuration-management procedure.

Firmware, AI model files, alarm thresholds, calibration parameters, communication settings, or other safety-related configuration items shall not be changed in the field without authorized engineering approval and a documented rollback method.

Any configuration change that could affect sensor acquisition, gas classification, alarm generation, fan operation, power consumption, communication behavior, diagnostic functions, or other characteristics relevant to the certified product shall be evaluated before release.

Following an approved firmware or configuration update, the GLD shall undergo the applicable verification procedure before being returned to service. This shall include confirmation of the installed version, sensor status, normal operating mode, alarm function, communication function, and absence of unresolved diagnostic faults.

Production and maintenance records shall identify the firmware version and, where applicable, the AI-model and configuration versions installed in each GLD unit.

g.18 Repair / Replacement Restrictions

Repair, replacement, and modification of the GLD shall be carried out only by trained and authorized personnel using approved procedures and approved replacement components.

No field modification shall be made to the enclosure, gas-sampling path, front stainless-steel mesh, cable-entry openings, antenna interface, alarm-module interface, grounding arrangement, enclosure joints, threads, sealing surfaces, or other Ex-critical construction features.

Drilling, welding, machining, re-threading, enlarging cable entries, bypassing protective devices, or any other unauthorized modification of the approved enclosure or protection-related construction is prohibited.

Replacement parts shall correspond to the approved part number, material, rating, manufacturer, and configuration identified in the controlled BOM or other approved technical documentation.

Substitution of the front mesh, fan, MQ sensor type, gasket, cable gland, antenna, alarm component, electrical protection device, enclosure part, or any other Ex-critical component shall not be permitted without controlled engineering review and, where applicable, certification assessment.

Damaged enclosure joints, threads, sealing surfaces, or other protection-related mechanical features shall not be repaired by uncontrolled re-machining, filing, grinding, coating, or other field rework.

Following repair or replacement, the affected assembly shall be inspected and the applicable functional and safety checks shall be completed before the GLD is returned to normal service.

The GLD shall be removed from service if visible damage, water ingress, an unresolved sensor fault, failed alarm verification, or loss of protective-earth continuity is identified.

g.19 Storage / Transport

During storage and transport, the GLD shall be protected against mechanical impact, excessive vibration, moisture ingress, corrosive contamination, dust accumulation, and other conditions that could damage the enclosure, sensing inlet, antenna, cable-entry components, grounding connection, or other controlled product features.

The detector shall be stored and transported in suitable protective packaging that prevents movement, impact, deformation, or abrasion of the enclosure and external components.

Particular care shall be taken to protect the front stainless-steel mesh, antenna, enclosure threads, sealing surfaces, cable-entry interfaces, and grounding stud from impact or contamination.

The GLD shall not be transported or stored with the gas-sampling inlet exposed to conditions likely to introduce excessive dust, oil, liquid, or other contaminants into the sensing area.

Covers, plugs, protective packaging, or other temporary transport protection used during shipping shall be removed before installation and shall not remain in place over the gas-sampling inlet during operation.

Before installation after storage or transport, the GLD shall be visually inspected for mechanical damage, corrosion, contamination, loose components, damaged seals, damaged mesh, or other abnormal conditions. Any damaged or suspect equipment shall be evaluated before being energized or placed into hazardous-area service.

The permitted storage-temperature range, storage humidity, packaging specification, and transportation environmental limits shall be defined in the final product specification and controlled logistics documentation.

g.20 Decommissioning

Decommissioning of the Gas Leak Detector (GLD) shall be performed only by trained and authorized personnel in accordance with the applicable site electrical-isolation and hazardous-area procedures.

Before removal of the equipment, the 24 VDC supply shall be isolated and verified de-energized. The applicable lockout/tagout procedure shall be applied where required.

The enclosure shall not be opened and electrical connections shall not be disconnected while an explosive atmosphere may be present unless the equipment has been de-energized and the area has been confirmed safe in accordance with the site procedure.

After the equipment has been safely isolated, the power wiring, grounding conductor, antenna connection, communication interfaces, mounting hardware, and associated field connections may be disconnected in accordance with the approved removal procedure.

Any cable entry, conduit, or field wiring remaining at the installation location after removal of the GLD shall be left in a condition that maintains the safety and integrity of the remaining installation. Open hazardous-area cable or conduit entries shall be closed or otherwise made safe in accordance with the applicable site requirements.

The removed GLD shall be identified as decommissioned and shall not be returned to hazardous-area service unless it has been inspected, repaired where necessary, and recommissioned in accordance with the approved procedures.

Equipment intended for disposal shall be handled in accordance with the applicable site and regulatory requirements for electronic equipment, metallic components, printed circuit boards, and other materials contained in the product.

Where the unit has been exposed to process gas, chemical contamination, or other hazardous substances, appropriate site decontamination and handling procedures shall be followed before repair, transport, storage, or disposal.

### 2.6.h · Nameplate Information (must include all ATEX-required marking information)

Nameplate artwork cannot yet be finalized: it depends on the certificate number, protection marking, temperature class, ambient range, IP rating, and serialization scheme, none of which has been assigned yet.

Status: Not yet available.

#### 2.6.h.1 Proposed nameplate content (draft — not a certified marking)

The nameplate artwork is finalized only after certification. The content below is the applicant's proposal for gas atmospheres only; no dust marking (category D / Ex t) is proposed, and the ambient range is the rated −20 °C to +60 °C.

| Field | Proposed content | Status |
| --- | --- | --- |
| Manufacturer | PT Galaksi Megatama Indonesia, Bekasi, Indonesia | Final |
| Product / model | Gas Leak Detector — GLD V2 | Final |
| Serial number | Per unit (sample: GLD2-0x1001) | Final |
| Year of manufacture | Per unit | Final |
| ATEX equipment group / category | II 2G | Proposed — pending ExCB |
| Ex marking | Ex db IIC T4 Gb | Proposed — pending ExCB |
| Ambient temperature | −20 °C ≤ Ta ≤ +60 °C | Final (rated range) |
| IP rating | IP66 | Proposed — test evidence pending |
| Electrical rating | 24 VDC, 8 W max. | Final |
| Cable entry | M20×1.5 | TBC |
| Certificate numbers | ATEX / IECEx certificate numbers | Pending certificate |
| CE marking + notified-body number | CE xxxx | Pending certificate |
| Warnings | “WARNING – DO NOT OPEN WHEN AN EXPLOSIVE ATMOSPHERE MAY BE PRESENT”; additional warnings (e.g. electrostatic charging) as required by the certificate | Proposed |

### 2.6.i · Ex Component Certificates

No components in this design currently hold an Ex component certificate. The processing unit (ESP32-S3-WROOM-1U-N16R8) and the communication module (E22-900MM22S) hold RF/EMC certifications (FCC, TELEC, CE, RoHS) — these are not Ex component certificates and do not satisfy this item.

Status: Not yet available.

#### 2.6.i.1 Components requiring Ex certificates

Ex-certified components needed for the proposed Ex d construction (gas group IIC). Each certificate must cover gas group IIC, the ambient range −20 °C to +60 °C and the thread form used.

| Component | Required certification | Supplier / certificate no. | Status |
| --- | --- | --- | --- |
| Flameproof enclosure (if a certified empty enclosure is used) | Ex db IIC component certificate (ATEX/IECEx) | TBC | Pending certificate |
| Cable gland, 24 VDC entry | Ex d barrier gland for IIC, M20×1.5, IP66 | TBC | Pending certificate |
| Blanking plug / adaptor | Ex d IIC, M20×1.5 | TBC | Pending certificate |
| Front mesh / flame arrestor (if claimed) | Certificate or test evidence of flame-arresting function | TBC | Pending certificate |
| Alarm beacon (if outside the flameproof enclosure) | Ex certificate of the beacon, or inclusion in the equipment assessment | TBC | Open |
| Antenna feed-through (if a certified bushing is used) | Ex d IIC bushing | TBC | Open |
| ESP32-S3 module, LoRa module | RF/EMC certificates only (FCC, TELEC, CE) — not Ex certificates | Espressif / Ebyte | N.A. for Ex |

# 3. Sample Information

## 3.1 · Model, Serial Number, and Status of the Sample

Sample submitted: model GLD V2 (GLD_V2), serial number GLD2-0x1001. The sample powers on and runs in Inference (normal operation) mode. Full sample register in Section 3.3.

Status: Available — sample identified (GLD2-0x1001); firmware/AI-model version to be recorded on the sample label.

## 3.2 · Necessary Test Fixtures or Auxiliary Equipment

Equipment anticipated for commissioning and bench-level functional verification:

| Equipment | Purpose |
| --- | --- |
| Regulated 24 VDC power supply | Provides a stable 24 VDC supply during commissioning and functional testing. |
| Digital multimeter | Measures actual voltage, verifies polarity and continuity, supports grounding/bonding verification. |
| Calibrated gas source / gas test kit | Provides the required gas types to verify correct sensor response. |
| Gas test chamber | Provides a controlled gas-exposure environment for reproducible test conditions. |
| Alarm load / relay simulator | Verifies that the alarm output and connected electrical load operate correctly. |
| LoRa peer / gateway | Verifies that data transmitted by the GLD is successfully received by the intended system. |
| Laptop and approved software | Reviews diagnostic information, configuration, logging, firmware identification; stores commissioning and test evidence. |
| Calibrated torque screwdriver | Ensures glands, terminals, enclosure fasteners, and grounding connections are tightened to approved torque values. |

Status: Draft plan only — not yet executed, and not evidence of ExCB/laboratory test readiness.

This is a working document, filled in against the internal working-draft template (IECEx ATEX Certification Information Requirements_Rev18092026.docx) supplied for this purpose. It follows that template's own section order rather than the more exhaustive item-by-item status tracking maintained in Dokumen_Teknis_Sertifikasi_GLD_IECEx_ATEX — the two documents are maintained in parallel and should stay consistent; treat the exhaustive tracking document as authoritative for item-level status where they differ. Data sources: the official product technical datasheet (Revision 4.0), internal technical specification documentation, EMC parameter measurement data, EasyEDA/JLCPCB design exports, and PT Galaksi Megatama Indonesia's legal documents.

#### 3.3 Sample register (current)

| Item | Value | Status |
| --- | --- | --- |
| Model | GLD V2 (GLD_V2) | Final |
| Serial number | GLD2-0x1001 | Final |
| Status | Powers on and runs in Inference / normal operation mode | Final |
| Hardware revision | Main board MotherBoardGLDVer2, schematic rev. 1.0 (19 Jul 2026) | Final |
| Firmware / AI model version | To be recorded on the sample label and test record | Open |
| Configuration | 24 VDC production configuration; battery path status per decision EX-17 | Decision required |

3.4 Test fixtures and auxiliary equipment (revision-guide list)

| Equipment | Purpose |
| --- | --- |
| Regulated 24 VDC power supply | Stable 24 VDC during commissioning and functional testing |
| Digital multimeter | Voltage, polarity, continuity and grounding/bonding verification |
| Calibrated gas source / gas test kit | Verify sensor response to the target gases |
| Gas test chamber | Reproducible controlled gas exposure |
| Alarm load / relay simulator | Verify the 24 V alarm output (J2) and its load |
| LoRa peer / gateway | Verify reception of transmitted data |
| Laptop and approved software | Diagnostics, configuration, logging, firmware identification |
| Calibrated torque screwdriver | Glands, terminals, enclosure fasteners and grounding to approved torque |

# Annex · Compliance matrix against the reviewer's revision guide

Each requirement of the revision guide, where it is addressed in this document, and its status.

| Guide item | Requirement | Section | Status |
| --- | --- | --- | --- |
| 1.1 | Application form | 1.1 | Final |
| 1.2 | Business licence / registration | 1.2 | Final |
| 1.3 | Organisation chart and contacts | 1.3 | Partial — QA/production functions to add |
| 1.4 | Plant address and facilities | 1.4 | Partial — PCB assembler to add |
| 1.5 | ISO 9001 | 1.5 | Open — audit planned Oct 2026 |
| 2.1–2.3 | Product description, specification, technical parameters | 2.1–2.3 | Final |
| 2.4 | Photo set (12 views) | 2.4, 2.4.1, 2.4.2 | Partial — 9/12 available; PCB underside, grounding-lug and antenna-bulkhead close-ups to add |
| 2.5 | Intended use and environment (IIC, T4, Zone 1, −20…+60 °C) | 2.5 | Final (proposed classification) |
| 6(a).1–6(a).6 | Assembly, section, enclosure, mesh, MQ arrangement, fan drawings | 2.6.a, 2.6.a.1 | Partial — controlled toleranced drawings to issue |
| 6(a).7 | Main PCB schematic | 2.6.a.7 | Final (main board); sensor module open |
| 6(a).8 | PCB layout | 2.6.a.8 | Partial — rev/date, laminate data, sensor module |
| 6(a).9–6(a).13 | Terminal, grounding, cable entry, antenna, alarm drawings | 2.6.a.1 (M-D14, M-D15, M-D5) | Partial — dimensioned drawings to issue |
| 6.b | Ex-critical BOM (EX-01…17, EL-01…13) | 2.6.b, 2.6.b.1–2 | Partial |
| 6.c | Non-metallic materials (MAT register) | 2.6.c, 2.6.c.1 | Open — supplier datasheets |
| 6.d | Manufacturing process (d.1–d.13) | 2.6.d, 2.6.d.1 | Partial — parameters TBC |
| 6.e | Ignition sources IS-01…17 and protection concept e.3 | 2.6.e, 2.6.e.1–3 | Partial |
| 6.f | Temperature class f.1–f.12 | 2.6.f, 2.6.f.1 | Partial — measurements outstanding |
| 6.g | Usage and installation instructions g.1–g.20 | 2.6.g, 2.6.g.1 | Draft available |
| 6.h | Nameplate (gas only, −20…+60 °C) | 2.6.h, 2.6.h.1 | Proposed |
| 6.i | Ex component certificates | 2.6.i, 2.6.i.1 | Pending certificate |
| 3.1–3.2 | Sample and fixtures | 3.1–3.4 | Final |